## Tabelle1
measurements with about 0.1mm Codewidth covered (initial measurements bevor Ringversuch 2019)
measurements with about 16mm code coverage, but only the central region evaluated for better comparability with the 0.1mm coverage measurements
Levelling staffs of "Ringversuch 2019"
Job 'n Staff					Basic Job Info					Tracewise Data Evaluation					Runwise Data Evaluation					Overall (Combined) Data Evaluation					Outliers and Reject Statistics					ISO 12858-1 Classifikation			Ambient Conditions				Additional Infos
Job Number	Staff Type	Staff Serial Number	Staff Owner	Sign Bug reprocessing	Calibration Date	Calibration Start	Calibration End	Duration [min]	Lowest Code-element [mm]	Highest Code-element [mm]	Num. of Runs	Pos. per Run	Number of Traces	Code-width covered [mm]	Pixel-Width [µm]	Center-Pixel	Lower Staff Support [mm]	Upper Staff Support [mm]	Zero Point Offset [mm]	Operator	Trace-wise Offset [µm]	Trace-wise STD Offset [µm]	Trace-wise Sigma Offset [µm]	Trace-wise Scale [ppm]	Trace-wise STD Scale [ppm]	Trace-wise Sigma Scale [ppm]	Trace-wise RMS [µm]	Tracewise Code-element Width Error [µm]	Tracewise Code-element Width STD [µm]	Tracewise Code-element Width Range [µm]	Runwise Offset [µm]	Runwise Sigma Offset [µm]	averaged runwise Offset [µm]	averaged Runwise Sigma Offset [µm]	Runwise Scale [ppm]	Runwise Sigma Scale [ppm]	averaged Runwise Scale [ppm]	averaged Runwise Sigma Scale [ppm]	Run-wise RMS [µm]	Overall Offset [µm]	Overall Sigma Offset [µm]	averaged Overall Offset [µm]	averaged Overall Sigma Offset [µm]	Overall Scale [ppm]	Overall Sigma Scale [ppm]	averaged Overall Scale [ppm]	averaged Overall Sigma Scale [ppm]	Overall RMS [µm]	Number of Measure-ments	Number of Outliers	Mean Edge Strength	Factor for removal of weak edges	Edge Strength STD	Number of Edge Strength Rejects	Mean Edge Noise	Factor for removal of noisy edges	Edge Noise STD	Number of Edge Noise Rejects	Mean Position-ing STD [µm]	Factor for Stability Removal	Overall Position-ing STD [µm]	Number of Position-ing Rejects	Tracewise ISO Class (individual)	Tracewise ISO Class (averaged)	ISO Class (Center Based)	average Tempe-rature [°C]	average Pressure [hPa]	average Humidity [%rH]	average CO2 Content [ppm]	Contractor	Contract Number	Remarks	Original File Name	Certificate Remarks
19	GPCL3	70554	swisstopo	TRUE	6/4/2019	13:32:51	16:05:24	152.6	55.6875	2968.6500	2	340	9	0.1	10.50	1224	470	2217		RP	-63.99	0.536	0.357	3.48	0.06	0.21	3.19	-29.53	2.334	6.34	-63.99	0.120	-63.929	0.322	3.48	0.069	3.432	0.185	3.23	-63.99	0.085	-63.959	0.227	3.48	0.049	3.443	0.131	3.25	6120	18	78.8	3	5.88	54	7.2	3	3.66	349	0.028	15	0.020	4	C	C	A	19.8	950.9	49.9	471.1	GSEG	-	still @ bugfixtesting for camera freeze, 0018 aborted due to E50003 "Controller ID mismatch" at Motorization??????	0019_20190604_1317_GPCL3_70554_swisstopo.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
29	GPCL3	70554	swisstopo	TRUE	6/7/2019	11:48:25	14:10:32	142.1	55.6875	2968.6500	2	340	9	0.1	10.50	1224	480	2210		RP	-63.37	0.690	0.334	3.77	0.21	0.19	2.99	-29.54	2.247	6.15	-63.38	0.114	-63.249	0.297	3.77	0.066	3.704	0.171	3.07	-63.38	0.087	-63.282	0.229	3.77	0.050	3.717	0.132	3.32	6120	16	73.2	3	5.78	56	6.3	3	3.28	325	0.032	15	0.020	2	C	C	A	19.9	950.4	49.8	428.5	GSEG	NA	measurement 0028 was aborted due to communications error with the motor controller. This is the repetition to 0019	0029_20190607_1130_GPCL3_70554_swisstopo.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
32	GPCL3	70554	swisstopo	TRUE	6/12/2019	10:03:42	12:25:19	141.6	55.6875	2968.6500	2	340	9	0.1	10.50	1224	420	2190		RP	-60.04	1.801	0.356	3.83	0.16	0.20	3.18	-30.38	2.502	6.83	-60.05	0.132	-59.977	0.314	3.83	0.076	3.780	0.180	3.55	-60.05	0.094	-59.978	0.225	3.83	0.054	3.779	0.129	3.58	6120	9	74.8	3	7.27	37	7.2	3	3.53	349	0.028	15	0.019	3	C	C	A	19.9	952.5	49.7	414.1	GSEG	Initial Test Phase		0032_20190612_0948_GPCL3_70554_swisstopo.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
110	Nedo GPCL3	070554 (swisstopo 34)	swisstopo	TRUE	11/12/2019	11:53:04	14:17:54	144.8	55.6875	2968.6500	2	340	171	15.9	10.41	1224	400	2195		RP	-62.58	6.064	0.503	3.52	3.54	0.29	4.52	-29.05	8.252	39.34	-62.20	0.064	-62.565	0.295	3.37	0.037	3.517	0.169	7.57	-62.20	0.046	-62.565	0.214	3.38	0.026	3.517	0.123	7.60	116280	622	75.7	3	8.84	615	7.6	3	5.00	5025	0.032	15	0.021	0	X	C	A	19.9	946.8	49.6	448.5	swisstopo	5703005198	lagerpunkt am oberen Lattenende sehr weit innen!	0110_20191112_1134_Nedo GPCL3_070554 (swisstopo 34)_swisstopo.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
127	NEDO GPCL3	070554 (swisstopo 34)	swisstopo	TRUE	11/20/2019	14:13:19	16:40:57	147.6	55.6875	2968.6500	2	340	171	15.9	10.39	1224	445	2215		RP	-62.46	5.369	0.490	3.13	3.50	0.28	4.40	-28.58	8.044	38.84	-62.47	0.065	-62.473	0.276	3.08	0.037	3.143	0.158	7.71	-62.46	0.046	-62.470	0.196	3.08	0.027	3.142	0.112	7.72	116280	845	70.4	3	8.40	549	7.0	3	5.33	4235	0.026	15	0.019	7	X	C	A	19.9	952.2	49.5	457.4	swisstopo	5703005198	Wiederholung zu Kal. 0110 (Kal.0124 & 0.125wegen Motorfehler abgebrochen, Kal 126 hat Kantenfehler (manuell entfernen), Latte nicht neu justiert)	0127_20191120_1358_NEDO GPCL3_070554 (swisstopo 34)_swisstopo.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
150	NEDO GPCL3	070554 (34)	swisstopo	TRUE	11/25/2019	14:38:59	17:03:06	144.1	55.6875	2968.6500	2	340	173	15.8	10.24	1224	400	2200		RP	-66.31	5.026	0.431	3.51	3.47	0.25	3.83	-29.38	8.438	40.63	-66.39	0.063	-66.383	0.160	3.51	0.036	3.536	0.092	7.37	-66.40	0.044	-66.388	0.114	3.51	0.025	3.538	0.065	7.37	117640	1338	74.0	3	9.19	513	7.5	3	4.90	5056	0.032	15	0.029	7	X	C	A	20.0	950.0	49.6	470.7	swisstopo	5703005198	0149 unterbrochen mit unbekanntem Fehler	0150_20191125_1424_NEDO GPCL3_070554 (34)_swisstopo.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
288	Leica GPCL3	070554 (swisstopo 34)	swisstopo	FALSE	4/21/2021	10:48:26	13:14:24	146.0	55.6875	2968.6500	2	340	173	15.9	10.29	1224	400	2200	0.000	RP	-59.31	4.834	0.434	3.48	3.50	0.25	3.81	-30.08	9.093	43.50	-59.49	0.060	-59.433	0.137	3.63	0.034	3.554	0.078	6.96	-59.49	0.042	-59.447	0.098	3.63	0.024	3.559	0.056	6.97	117640	1151	71.9	3	9.81	166	8.2	3	3.55	9785	0.064	15	0.213	0	X	C	A	20.1	955.6	49.8	446.9	swisstopo	NP Experiment		0288_20210421_1036_Leica GPCL3_070554 (swisstopo 34)_swisstopo.mat	none
329	Leica GPCL3	70554 (swisstopo 34)	swisstopo	FALSE	3/17/2022	9:53:30	12:19:33	146.0	55.6875	2968.6500	2	340	173	15.9	10.30	1224	400	2200	0.000	RP	-64.82	5.216	0.424	3.25	3.45	0.24	3.76	-29.33	8.862	41.82	-64.91	0.067	-64.873	0.159	3.29	0.038	3.275	0.091	7.80	-64.91	0.047	-64.873	0.112	3.29	0.027	3.275	0.064	7.81	117640	1375	70.9	3	8.58	273	7.7	3	4.87	4879	0.045	15	0.039	20	X	C	B	19.9	964.6	48.5	479.1	swisstopo	5703006202		0329_20220317_0941_Leica GPCL3_70554 (swisstopo 34)_swisstopo.mat	none
344	Leica GPCL3	70554 (swisstopo 34)	swisstopo	FALSE	3/24/2023	14:25:44	16:46:26	140.7	55.6875	2968.6500	2	340	173	15.9	10.28	1224	400	2200	0.000	RoPr	-73.86	5.050	0.413	3.45	3.41	0.24	3.68	-29.66	8.323	38.79	-73.91	0.066	-73.910	0.145	3.38	0.038	3.476	0.083	7.74	-73.91	0.047	-73.910	0.104	3.39	0.027	3.475	0.059	7.75	117640	1181	67.3	3	7.28	759	6.1	3	4.37	4898	0.045	15	0.036	11	X	C	A	20.0	952.0	49.4	546.5	swisstopo	5703006666		0344_20230324_1412_Leica GPCL3_70554 (swisstopo 34)_swisstopo.mat	none
345	Leica GPCL3	70554 (swisstopo 34)	swisstopo	FALSE	3/27/2023	9:29:28	11:51:05	141.6	55.6875	2968.6500	2	340	173	15.9	10.28	1224	510	2500	0.000	RoPr	-67.31	4.870	0.420	3.39	3.41	0.24	3.72	-29.63	8.333	38.70	-67.49	0.060	-67.343	0.160	3.53	0.034	3.397	0.092	6.99	-67.50	0.043	-67.344	0.124	3.53	0.024	3.397	0.071	7.03	117640	1204	67.4	3	7.27	766	6.1	3	4.39	4954	0.046	15	0.031	17	X	C	A	20.1	956.9	49.4	430.6	swisstopo	5703006666		0345_20230327_0911_Leica GPCL3_70554 (swisstopo 34)_swisstopo.mat	none
369	Leica GPCL3	70554 (swisstopo 34)	swisstopo	FALSE	2/29/2024	10:50:41	13:45:35	174.9	55.6875	2968.6500	2	340	173	15.9	10.27	1224	510	2500	0.000	RoPr	-65.10	6.110	0.408	3.42	3.53	0.23	3.65	-29.98	8.570	40.39	-65.21	0.063	-65.138	0.146	3.51	0.036	3.438	0.084	7.49	-65.21	0.045	-65.133	0.125	3.50	0.026	3.436	0.072	7.55	117640	1159	74.2	3	8.10	699	7.1	3	4.83	5033	0.067	15	0.043	3	X	C	A	20.0	953.3	49.4	500.0	swisstopo	5703006666 (2024)		0369_20240229_1034_Leica GPCL3_70554 (swisstopo 34)_swisstopo.mat	none
370	Leica GPCL3	70554 (swisstopo 34)	swisstopo	FALSE	2/29/2024	15:28:15	18:22:17	174.0	55.6875	2968.6500	2	340	173	15.9	10.26	1224	510	2500	0.000	RoPr	-71.88	5.468	0.407	3.45	3.52	0.23	3.65	-29.98	8.527	40.12	-72.04	0.059	-71.925	0.147	3.61	0.034	3.474	0.084	6.94	-72.03	0.042	-71.916	0.112	3.59	0.024	3.467	0.064	6.97	117640	1204	70.0	3	7.65	678	6.6	3	4.57	5238	0.060	15	0.043	1	X	C	A	20.0	949.5	49.4	531.3	swisstopo	5703006666 (2024)	Wiederholung zu 0369	0370_20240229_1515_Leica GPCL3_70554 (swisstopo 34)_swisstopo.mat	none
	Mittelwert					-65.13	±4.180			3.50	±0.196

11	GPCL3	26097	GSEG	TRUE	5/31/2019	14:12:28	16:32:38	140.2	55.6875	2968.6500	2	340	9	0.1	10.50	1224	545	2370		RP	48.55	1.647	0.282	-5.24	0.18	0.16	2.56	37.75	1.531	4.29	48.55	0.105	48.491	0.268	-5.23	0.060	-5.205	0.155	2.85	48.56	0.074	48.490	0.189	-5.24	0.043	-5.205	0.109	2.85	6120	82	62.8	3	3.20	60	8.8	3	4.31	20	0.034	15	0.030	0	X	X	A	20.0	964.5	49.4	478.5	GSEG	NA	Levelling Rod cleaned (Water only!)	0011_20190531_1356_GPCL3_26097.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
31	GPCL3	26097	GSEG	TRUE	6/11/2019	13:28:07	15:49:45	141.6	55.6875	2968.6500	2	340	9	0.1	10.50	1224	475	2255		RP	39.61	0.518	0.316	-4.75	0.07	0.18	2.86	38.78	1.630	4.58	39.61	0.106	39.653	0.305	-4.76	0.061	-4.746	0.176	2.88	39.62	0.080	39.660	0.230	-4.76	0.046	-4.749	0.133	3.07	6120	83	62.6	3	3.40	63	9.1	3	4.29	17	0.034	15	0.023	7	X	X	A	19.8	951.5	49.8	410.5	GSEG	NA	measurement 0030 interrupted (motor error), this is the repetition to 0011	0031_20190611_1312_GPCL3_26097_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
39	GPCL3	26097	GSEG	TRUE	6/13/2019	9:39:12	12:01:53	142.7	55.6875	2968.6500	2	340	9	0.1	10.50	1224	535	2285		RP	44.92	0.615	0.275	-5.15	0.20	0.16	2.47	37.88	1.654	4.67	44.87	0.092	44.981	0.262	-5.12	0.053	-5.173	0.151	2.47	44.86	0.065	44.947	0.185	-5.12	0.037	-5.157	0.106	2.48	6120	80	61.0	3	3.24	48	9.3	3	4.25	36	0.035	15	0.023	8	X	C	A	20.0	955.8	49.7	421.5	GSEG	Initial Test Phase	Latte im Bereich unmittelbar nach dem Fuss stark verzogen ... Lattenmitte wandert um ca. 20Px (ca. 0.2mm) nach links aus. Messung 0034 abgegrochen, Kommunikationsfehler mit Motorcontroller / Motorcontroller, Messungen 0035 bis 038 gleicher Seich	0039_20190613_0921_GPCL3_26097_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
73	Leica GPCL3	26097	GSEG	TRUE	10/7/2019	9:47:56	12:16:27	148.5	55.6875	2968.6500	2	340	167	15.9	10.62	1224	420	2200		RP	53.38	4.144	0.320	-4.86	0.84	0.19	2.88	48.59	8.638	40.58	53.43	0.044	53.438	0.172	-4.95	0.025	-4.916	0.099	5.13	53.43	0.031	53.424	0.124	-4.95	0.018	-4.902	0.071	5.13	113560	1538	61.8	3	3.78	2199	9.0	3	4.29	432	0.034	15	0.025	0	X	X	A	19.8	958.5	49.8	415.3	GSEG	Initial Test Phase		0073_20191007_0927_Leica GPCL3_26097_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
92	Leica GPCL3	26097	GSEG	TRUE	10/11/2019	11:57:52	14:23:03	145.2	55.6875	2968.6500	2	340	169	15.9	10.50	1224	420	2200		RP	54.33	5.232	0.315	-5.13	0.56	0.18	2.84	48.48	8.645	40.37	54.20	0.049	54.283	0.193	-5.12	0.028	-5.093	0.111	5.75	54.21	0.035	54.284	0.140	-5.12	0.020	-5.093	0.081	5.77	114920	1437	61.8	3	3.58	2226	8.5	3	4.30	436	0.031	15	0.024	1	X	X	A	19.7	960.8	49.9	455.9	GSEG	Initial Test Phase		0092_20191011_1142_Leica GPCL3_26097_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
221	Leica GPCL3	26097	GSEG	TRUE	6/11/2020	14:25:07	16:51:29	146.4	55.6875	2968.6500	2	340	173	15.9	10.29	1224	420	2200	0.000	RP	42.62	5.514	0.290	-4.40	1.04	0.17	2.61	48.24	8.678	40.78	42.42	0.049	42.430	0.146	-4.34	0.028	-4.302	0.084	5.79	42.43	0.035	42.434	0.109	-4.35	0.020	-4.305	0.062	5.81	117640	2196	60.0	3	3.90	2011	8.3	3	4.33	826	0.055	15	0.114	7	X	X	A	19.9	948.4	50.4	396.0	GSEG	Vorzeichentest	Interfero - Cal (schlechte Latte)	0221_20200611_1404_Leica GPCL3_26097_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Test with (Interfero - Cal), "regular Targetpositions"
224	Leica GPCL3	26097	GSEG	FALSE	6/15/2020	11:13:32	13:43:49	150.3	55.6875	2968.6500	2	340	173	15.9	10.29	1224	420	2200	0.000	RP	51.54	5.034	0.291	-4.10	1.04	0.17	2.62	48.18	8.682	40.89	51.52	0.050	51.321	0.147	-4.07	0.029	-3.972	0.084	5.91	51.53	0.035	51.339	0.110	-4.07	0.020	-3.982	0.063	5.93	117640	2638	64.3	3	4.26	1922	9.3	3	4.63	796	0.056	15	0.130	2	X	X	A	19.9	958.7	50.5	410.9	GSEG	Vorzeichentest	Interfero + Cal (schlechte Latte)	0224_20200615_1102_Leica GPCL3_26097_GSEG.mat	Test with (Interfero + Cal), "regular Targetpositions"
226	Leica GPCL3	26097	GSEG	FALSE	6/16/2020	11:31:19	14:04:32	153.2	55.6875	2968.6500	2	340	173	15.9	10.29	1224	420	2200	0.000	RP	57.29	5.062	0.294	-4.31	1.01	0.17	2.65	48.20	8.684	40.73	57.34	0.049	57.138	0.154	-4.28	0.028	-4.203	0.088	5.84	57.34	0.035	57.119	0.123	-4.28	0.020	-4.196	0.070	5.90	117640	2453	60.6	3	3.94	2047	8.6	3	4.34	1157	0.056	15	0.123	2	X	X	A	19.9	956.5	50.7	428.5	GSEG	Vorzeichentest	Target Positions scaled with the Results of Calibration 0224 to minimize the distance from the Camera Center and test the results. (Interfero + Cal)	0226_20200616_1117_Leica GPCL3_26097_GSEG.mat	Test with (Interfero + Cal), Targetpositions scaled and shifted with parameters found in Cal.0224
228	Leica GPCL3	26097	GSEG	FALSE	6/16/2020	16:28:49	18:56:57	148.1	55.6875	2968.6500	2	340	173	15.9	10.29	1224	420	2200	0.000	RP	52.13	5.030	0.298	-4.27	1.01	0.17	2.64	48.20	8.682	40.82	52.16	0.050	51.891	0.153	-4.23	0.029	-4.131	0.087	5.80	52.16	0.035	51.932	0.123	-4.23	0.020	-4.150	0.070	5.85	117640	2097	61.3	3	3.97	2089	8.7	3	4.34	1190	0.045	15	0.032	26	X	X	A	19.8	954.4	50.4	459.6	GSEG	Vorzeichentest	Target Positions scaled with the Results of Calibration 0224 + added noise (±0.2mm) to compensate for scale and offset but still deviate from the "true" edge positions. (Interfero + Cal)	0228_20200616_1616_Leica GPCL3_26097_GSEG.mat	Test with (Interfero + Cal), Targetpositions scaled and shifted with paramters of 0224 + random offset
229	Leica GPCL3	26097	GSEG	TRUE	6/17/2020	8:36:48	11:08:15	151.5	55.6875	2968.6500	2	340	173	15.9	10.29	1224	420	2200	0.000	RP	50.60	5.049	0.293	-4.19	0.98	0.17	2.63	48.10	8.648	40.65	50.61	0.050	50.432	0.151	-4.16	0.029	-4.075	0.087	5.88	50.61	0.035	50.485	0.125	-4.15	0.021	-4.100	0.072	5.94	117640	2163	59.9	3	3.83	2216	8.5	3	4.32	1223	0.052	15	0.086	7	X	X	A	19.8	952.0	50.7	411.4	GSEG	Vorzeichentest	Test with "bad staff", nominal postions scaled and offseted with the parameter found in Calibration 0224 to come a sclose to the "true" edge positions as possible. (Interfero - Cal)	0229_20200617_0819_Leica GPCL3_26097_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Test with (Interfero - Cal), TargetPositions scaled and shifted with parameters found in Cal. 0224
230	Leica GPCL3	26097	GSEG	TRUE	6/17/2020	12:17:14	14:47:13	150.0	55.6875	2968.6500	2	340	173	15.9	10.29	1224	420	2200	0.000	RP	52.08	5.071	0.295	-4.55	0.97	0.17	2.64	48.24	8.669	40.72	52.12	0.049	51.872	0.150	-4.50	0.028	-4.423	0.086	5.81	52.13	0.035	51.981	0.124	-4.50	0.020	-4.465	0.071	5.86	117640	2162	59.9	3	3.83	2225	8.4	3	4.31	1225	0.048	15	0.054	7	X	X	A	19.8	953.1	50.7	403.1	GSEG	Vorzeichentest	Test with "bad staff", nominal postions scaled and offseted with the parameter found in Calibration 0224 plus a random offset for each target position (±0.2mm). (Interfero - Cal)	0230_20200617_1203_Leica GPCL3_26097_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Test with (Interfero - Cal), Targetpositions scaled and shifted with parameters determined in Cal.0224 + random offset at each position
232	Leica GPCL3	26097	GSEG	FALSE	6/22/2020	14:33:04	17:02:30	149.4	55.6875	2968.6500	2	340	173	15.9	10.30	1224	420	2220	0.000	RP	55.23	5.139	0.310	-3.95	0.50	0.18	2.80	48.40	8.663	40.85	55.22	0.048	55.158	0.165	-3.96	0.028	-3.933	0.094	5.69	55.22	0.034	55.137	0.117	-3.96	0.020	-3.923	0.067	5.69	117640	2000	60.8	3	3.87	2289	8.5	3	4.35	945	0.079	15	0.267	0	X	X	A	19.6	963.9	46.9	431.8	GSEG	Vorzeichentest	Test with 4 meteosensors at 100/1150/2316/2930mm. Computer and all other possible heatsources removed further away from the rod. (Interfero + Cal)	0232_20200622_1421_Leica GPCL3_26097_GSEG.mat	Test with unmodified scale additional meteosensors for staff temperature measurement added
	Mittelwert	26097					50.17	±5.128			-4.56	±0.436

234	Leica GPCL3	26097	GSEG	FALSE	6/23/2020	11:53:23	14:22:43	149.3	55.6875	2968.6500	2	340	173	15.9	10.30	1224	420	2200	0.000	RP	51.01	5.324	0.295	0.43	0.54	0.17	2.66	48.30	8.671	40.77	51.00	0.049	50.854	0.149	0.46	0.028	0.515	0.085	5.79	51.01	0.037	50.931	0.193	0.46	0.021	0.481	0.110	6.17	117640	2189	60.6	3	3.89	1780	8.5	3	4.35	934	0.056	15	0.101	7	X	X	A	19.7	962.7	53.0	465.1	GSEG	Vorzeichentest	Test with modified Staff 26097, the tension spring for the invar tape has been tightened (moved upwards) for about 4mm increasing the tension applied to the invar tape which should result in a higher scale. Additional meteo-sensors attached to the staff for measuring material temperatue (M01@100mm, S01@1033mm, S02@1966mm, S03@2900mm, meteo measurement started at 11:16:30). (Interfero + Cal)	0234_20200623_1138_Leica GPCL3_26097_GSEG.mat	Test with higher tension of invar tape (spring appr. 4mm higher tension) in order to increase scale additional meteo sensors for staff temperature measurement (M01@100m/S01@1033mm/S02@1966mm/S03@2900mm)
236	Leica GPCL3	26097	GSEG	FALSE	6/23/2020	18:15:47	20:46:29	150.7	55.6875	2968.6500	2	340	173	15.9	10.30	1224	420	2200	0.000	RP	53.89	5.157	0.307	0.23	0.56	0.18	2.73	48.47	8.676	41.05	53.90	0.048	53.733	0.146	0.22	0.028	0.290	0.083	5.67	53.91	0.034	53.702	0.104	0.21	0.020	0.299	0.059	5.68	117640	2918	60.6	3	4.17	1447	8.6	3	4.66	870	0.049	15	0.055	11	X	X	A	19.7	961.4	50.2	424.6	GSEG	Vorzeichentest	Wiederholung von Test 0234	0236_20200623_1804_Leica GPCL3_26097_GSEG.mat	Repetition to measurement 0234 which showed a big difference between the two individual runs.
240

	Mittelwert	26097					52.46	±1.448			0.33	±0.122

20	GPCL2	75177	GSEG	TRUE	6/5/2019	9:15:31	10:45:15	89.7	55.6875	1930.8375	2	215	9	0.1	10.50	1224	339	1740		RP	-20.97	0.404	0.376	3.76	0.10	0.34	2.75	-14.12	2.360	6.45	-20.96	0.126	-20.850	0.343	3.76	0.113	3.660	0.309	2.78	-20.96	0.089	-20.848	0.243	3.76	0.080	3.659	0.219	2.79	3870	24	70.3	3	5.00	10	8.1	3	4.10	74	0.014	15	0.011	4	C	C	A	19.8	950.0	50.0	456.7	GSEG	NA		0020_20190605_0859_GPCL2_75177_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
22	GPCL2	75177	GSEG	TRUE	6/5/2019	13:30:54	15:00:27	89.5	55.6875	1930.8375	2	215	9	0.1	10.50	1224	343	1765		RP	-14.39	0.742	0.364	3.15	0.21	0.33	2.69	-14.51	2.254	6.21	-14.38	0.123	-14.330	0.334	3.17	0.111	3.120	0.301	2.74	-14.38	0.087	-14.331	0.236	3.17	0.078	3.120	0.212	2.74	3870	22	68.0	3	4.88	2	7.6	3	3.90	87	0.016	15	0.010	2	C	C	A	19.8	947.0	49.8	456.7	GSEG	NA	Repetition to 0020.	0022_20190605_1315_GPCL2_75177_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
24	GPCL2	75177	GSEG	TRUE	6/6/2019	8:55:51	10:26:08	90.3	55.6875	1930.8375	2	215	9	0.1	10.50	1224	320	1740		RP	-19.38	0.820	0.368	3.72	0.13	0.33	2.70	-14.34	2.387	6.53	-19.39	0.124	-19.361	0.332	3.76	0.112	3.713	0.299	2.74	-19.39	0.090	-19.368	0.241	3.75	0.081	3.720	0.217	2.82	3870	14	68.1	3	4.87	17	7.7	3	4.00	95	0.013	15	0.007	7	C	C	A	19.7	954.5	49.9	415.3	GSEG	NA	Repetition to 0020 & 0022	0024_20190606_0839_GPCL2_75177_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
83	Leica GPCL2	75177	GSEG	TRUE	10/9/2019	13:34:48	15:06:09	91.4	55.6875	1934.8875	2	216	167	15.9	10.61	1224	340	1750		RP	-8.33	5.868	0.461	2.56	1.47	0.41	3.37	-14.43	9.086	39.41	-8.65	0.064	-8.381	0.303	2.56	0.058	2.597	0.273	6.06	-8.65	0.046	-8.381	0.217	2.55	0.041	2.597	0.195	6.07	72144	311	65.8	3	5.33	737	7.7	3	3.89	1337	0.021	15	0.016	0	X	B	A	19.8	955.3	49.8	435.0	GSEG	Inital Test Phase		0083_20191009_1319_Leica GPCL2_75177_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
85	Leica GPCL2	75177	GSEG	TRUE	10/10/2019	8:30:05	10:03:19	93.2	55.6875	1934.8875	2	216	167	15.9	10.62	1224	340	1750		RP	-13.55	5.028	0.462	2.54	1.18	0.41	3.38	-14.49	9.079	39.00	-13.60	0.061	-13.576	0.304	2.53	0.055	2.559	0.274	5.82	-13.60	0.043	-13.576	0.219	2.53	0.039	2.559	0.197	5.83	72144	347	69.7	3	5.63	763	8.4	3	4.07	1351	0.027	15	0.027	0	X	B	A	19.7	958.3	49.9	412.3	GSEG	Inital Test Phase		0085_20191010_0813_Leica GPCL2_75177_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	75177					-15.40	±4.397			3.15	±0.542

21	GPCL2	75170	GSEG	TRUE	6/5/2019	11:34:06	13:03:06	89.0	55.6875	1930.8375	2	215	9	0.1	10.50	1224	310	1770		RP	-4.66	1.345	0.352	2.35	0.21	0.32	2.58	-9.30	2.065	5.71	-4.66	0.129	-4.674	0.323	2.36	0.117	2.361	0.292	2.86	-4.65	0.096	-4.664	0.244	2.33	0.086	2.354	0.220	2.99	3870	22	69.6	3	5.12	26	7.9	3	4.09	46	0.014	15	0.011	2	C	B	A	19.8	948.4	50.0	459.6	GSEG	NA		0021_20190605_1115_GPCL2_75170_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
23	GPCL2	75170	GSEG	TRUE	6/5/2019	15:49:16	17:18:38	89.4	55.6875	1930.8375	2	215	9	0.1	10.50	1224	333	1755		RP	-6.13	1.165	0.362	2.49	0.06	0.33	2.63	-9.17	2.012	5.58	-6.14	0.130	-6.182	0.334	2.50	0.117	2.525	0.302	2.85	-6.15	0.092	-6.183	0.236	2.50	0.083	2.526	0.213	2.84	3870	33	67.4	3	4.80	26	7.5	3	3.91	59	0.016	15	0.012	7	C	B	A	19.7	945.1	49.7	451.0	GSEG	NA	Repetition to 0021.	0023_20190605_1527_GPCL2_75170_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
25	GPCL2	75170	GSEG	TRUE	6/6/2019	11:02:05	12:32:04	90.0	55.6875	1930.8375	2	215	9	0.1	10.50	1224	320	1740		RP	-12.34	1.606	0.354	2.66	0.11	0.32	2.58	-9.42	2.057	5.68	-12.37	0.137	-12.360	0.330	2.69	0.123	2.684	0.297	3.00	-12.37	0.097	-12.361	0.233	2.69	0.087	2.684	0.210	3.00	3870	28	68.2	3	4.95	27	7.5	3	3.97	46	0.014	15	0.009	4	C	B	A	19.7	955.1	49.8	430.9	GSEG	NA	Repetition to 0021 & 0023	0025_20190606_1046_GPCL2_75170_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
84	Leica GPCL2	75170	GSEG	TRUE	10/9/2019	15:43:29	17:16:55	93.4	55.6875	1934.8875	2	216	167	15.8	10.61	1224	320	1750		RP	-3.32	4.889	0.396	2.34	3.79	0.36	2.92	-5.47	9.060	37.95	-3.42	0.058	-3.345	0.281	2.54	0.053	2.373	0.253	5.56	-3.42	0.041	-3.345	0.199	2.54	0.037	2.373	0.179	5.57	72144	417	64.3	3	5.66	593	7.6	3	3.89	1079	0.022	15	0.017	0	C	A	A	19.8	955.9	49.8	431.4	GSEG	Inital Test Phase		0084_20191009_1527_Leica GPCL2_75170_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
86	Leica GPCL2	75170	GSEG	TRUE	10/10/2019	10:31:43	12:04:28	92.7	55.6875	1934.8875	2	216	167	15.8	10.61	1224	320	1750		RP	1.60	5.110	0.401	2.27	3.83	0.36	2.94	-5.45	9.052	37.89	1.56	0.061	1.595	0.283	2.37	0.055	2.288	0.255	5.78	1.55	0.043	1.590	0.203	2.38	0.039	2.293	0.183	5.79	72144	436	65.0	3	5.73	619	7.7	3	3.93	1075	0.022	15	0.018	4	C	A	A	19.8	959.8	49.9	427.6	GSEG	Inital Test Phase		0086_20191010_1015_Leica GPCL2_75170_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	75170					-5.01	±4.498			2.49	±0.128

TUM 2019	GPCL 3	26090	GSEG		6/25/2019					5.00				-5.02					A
41	GPCL 3	26090	GSEG	TRUE	6/14/2019	10:48:35	13:06:40	138.1	55.6875	2968.6500	2	340	9	0.1	10.50	1224	615	2265		RP	39.88	2.111	0.258	-3.95	0.18	0.15	2.36	46.29	1.828	5.26	39.94	0.106	39.856	0.239	-3.98	0.061	-3.934	0.137	2.92	39.95	0.076	39.855	0.171	-3.98	0.043	-3.934	0.098	2.94	6120	57	58.6	3	3.16	61	7.1	3	3.96	19	0.036	15	0.022	1	X	X	A	19.9	957.9	49.8	451.5	GSEG	Inital Test Phase	Meaurement 0040 was aborted as one of the edges is badly damaged. This caused an unintended abortion of the whole calibration. This is now corrected. Before this measurement the motorcontroller was changed to a new model. Hope all the interruptions are gone!!!	0041_20190614_1033_GPCL 3_26090_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
53	GPCL3	26090	GSEG	TRUE	6/19/2019	10:38:16	13:02:53	144.6	55.6875	2968.6500	2	340	9	0.1	10.50	1224	585	2195		RP	48.29	1.085	0.254	-2.91	0.07	0.15	2.31	44.73	1.725	4.92	48.30	0.092	48.249	0.234	-2.92	0.053	-2.873	0.134	2.50	48.29	0.066	48.231	0.171	-2.91	0.038	-2.856	0.098	2.55	6120	61	62.7	3	3.27	65	8.1	3	4.07	27	0.040	15	0.026	7	X	X	A	19.8	953.4	49.6	432.1	GSEG	Inital Test Phase	Measurement 0051 was aborted due to a timeout from the camera :( / Measurement 0052 was aborted with the drive rod outside the drive whell (position -440mm), with no idea what the reason could be.	0053_20190619_1019_GPCL3_26090_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
56	GPCL3	26090	GSEG	TRUE	6/20/2019	9:08:01	11:31:49	143.8	55.6875	2968.6500	2	340	9	0.1	10.50	1224	535	2190		RP	48.19	1.293	0.254	-3.18	0.11	0.15	2.31	44.45	1.724	4.94	48.20	0.094	48.150	0.234	-3.19	0.054	-3.149	0.134	2.58	48.20	0.067	48.190	0.166	-3.19	0.038	-3.168	0.095	2.59	6120	75	61.3	3	3.24	62	7.8	3	4.08	17	0.041	15	0.027	9	X	X	A	19.8	955.5	49.7	415.4	GSEG	Initial Test Phase	with meas 0055 the camera calibration failed (Residuals upt o 250µm)	0056_20190620_0851_GPCL3_26090_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
65	GPCL3	26090	GSEG	TRUE	7/10/2019	8:43:45	11:10:16	146.5	55.6875	2968.6500	2	340	169	15.9	10.50	1224	535	2190		RP	47.31	6.783	0.355	-3.72	0.76	0.20	3.23	54.87	13.608	57.37	47.29	0.061	47.298	0.183	-3.72	0.035	-3.715	0.105	7.30	47.28	0.043	47.353	0.129	-3.72	0.025	-3.741	0.074	7.30	114920	931	63.0	3	4.03	1007	7.9	3	4.19	467	0.038	15	0.025	0	X	X	A	19.7	958.9	49.7	396.8	GSEG	Inital TestPhase after Ringexperiment	Repetition to 0064 which was aborted due to a matlab indexing error	0065_20190710_0825_GPCL3_26090_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
75	Leica GPCL3	26090	GSEG	TRUE	10/7/2019	16:06:16	18:33:33	147.3	55.6875	2968.6500	2	340	167	15.9	10.62	1224	650	2250		RP	51.37	6.933	0.353	-4.21	0.72	0.20	3.21	54.94	13.662	57.39	51.35	0.063	51.387	0.174	-4.20	0.036	-4.223	0.100	7.45	51.35	0.045	51.359	0.124	-4.20	0.026	-4.209	0.071	7.45	113560	995	61.7	3	3.84	1020	7.7	3	4.10	405	0.036	15	0.023	0	X	X	A	19.8	958.9	49.8	423.3	GSEG	Initial Test Phase	0074 abgebrochen wegen interferometerfehler	0075_20191007_1545_Leica GPCL3_26090_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
195	Leica GPCL3	26090	GSEG	TRUE	2/27/2020	9:02:09	11:28:46	146.6	55.6875	2968.6500	2	340	463	18.9	10.24	1224	650	2250	0.000	RP	48.17	8.359	0.360	-4.07	1.08	0.21	3.27	53.52	13.212	60.83	48.40	0.044	48.208	0.189	-4.03	0.025	-4.037	0.108	8.51	48.40	0.031	48.206	0.141	-4.03	0.018	-4.037	0.081	8.53	314840	3827	61.1	3	4.39	1347	7.4	3	4.33	3045	0.083	15	0.075	10	X	X	A	20.2	946.3	49.1	419.3	GSEG	Nedo Test	Wie Versuche 0191 bis 01914 mit SN76037. Latte 26090 zeigt eine andere Kantencharakteristik. CodeCovaerage 19mm, TraceDistance 0.05mm	0195_20200227_0848_Leica GPCL3_26090_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	26090					47.24	±3.497			-3.67	±0.470

42	GPCL 3	26089	GSEG	TRUE	6/14/2019	13:50:49	16:09:07	138.3	55.6875	2968.6500	2	340	9	0.1	10.50	1224	615	2300		RP	41.04	0.856	0.371	-2.96	0.21	0.21	3.42	26.93	1.983	5.65	41.01	0.125	41.073	0.358	-2.95	0.072	-2.973	0.206	3.45	41.03	0.091	41.073	0.262	-2.96	0.053	-2.973	0.150	3.58	6120	10	61.3	3	3.10	36	8.0	3	4.21	12	0.043	15	0.029	0	C	C	A	19.8	956.2	49.8	448.9	GSEG	Inital Test Phase		0042_20190614_1334_GPCL 3_26089_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
60	GPCL3	26089	GSEG	TRUE	6/21/2019	9:18:38	11:42:29	143.8	55.6875	2968.6500	2	340	9	0.1	10.50	1224	620	2340		RP	45.95	1.050	0.373	-2.98	0.15	0.21	3.42	21.36	1.972	5.60	45.95	0.130	45.972	0.355	-2.98	0.075	-3.002	0.204	3.58	45.95	0.092	45.972	0.252	-2.98	0.053	-3.002	0.144	3.59	6120	4	57.6	3	2.99	41	7.1	3	4.23	36	0.034	15	0.023	0	C	C	A	19.7	958.1	50.0	441.0	GSEG	Inital Testphase		0060_20190621_0901_GPCL3_26089_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
70	Leica GPCL3	26089	GSEG	TRUE	10/3/2019	14:58:00	17:25:55	147.9	55.6875	2968.6500	2	340	167	15.9	10.62	1224	620	2340		RP	52.53	5.475	0.318	-2.24	1.36	0.18	2.89	36.06	11.743	55.43	52.46	0.050	52.519	0.185	-2.21	0.029	-2.221	0.106	5.94	52.46	0.036	52.518	0.136	-2.21	0.021	-2.221	0.078	5.96	113560	883	56.4	3	3.32	1655	6.8	3	4.04	819	0.033	15	0.020	1	X	C	A	19.9	957.7	49.5	433.0	GSEG	Initial Test Phase		0070_20191003_1441_Leica GPCL3_26089_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
91	Leica GPCL3	26089	GSEG	TRUE	10/11/2019	9:06:51	11:33:53	147.0	55.6875	2968.6500	2	340	169	15.9	10.50	1224	620	2340		RP	52.31	5.556	0.347	-3.07	1.75	0.20	3.14	35.92	11.743	55.25	52.25	0.052	52.298	0.213	-3.04	0.030	-3.051	0.122	6.17	52.26	0.037	52.300	0.158	-3.04	0.021	-3.052	0.091	6.20	114920	812	60.6	3	3.74	1636	7.4	3	4.28	859	0.030	15	0.020	8	X	C	A	19.7	962.1	49.7	452.0	GSEG	Initial Test Phase	Camera lowered for about 10mm, whloe setup newly adjusted	0091_20191011_0851_Leica GPCL3_26089_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	26089					47.92	±4.763			-2.97	±0.464

TUM 2019	GPCL3	26088	GSEG					22.00				-5.90
44	GPCL3	26088	GSEG	TRUE	6/17/2019	8:49:22	11:08:23	139.0	55.6875	2968.6500	2	340	9	0.1	10.50	1224	590	2295		RP	26.70	0.285	0.285	-4.10	0.17	0.16	2.62	39.22	1.676	4.72	26.71	0.095	26.677	0.270	-4.11	0.055	-4.085	0.155	2.63	26.71	0.068	26.697	0.192	-4.11	0.039	-4.091	0.110	2.66	6120	16	63.5	3	3.29	49	10.5	3	4.95	30	0.037	15	0.023	1	X	X	A	19.9	959.4	49.8	409.3	GSEG	Inital Test Phase	Measurement 0043 interrupted due to automatic windows update on GSEG-Meteo. 26088 is missing the cover for the tension spring on top of the rod	0044_20190617_0834_GPCL3_26088_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
61	GPCL3	26088	GSEG	TRUE	6/21/2019	12:18:16	14:42:13	144.0	55.6875	2968.6500	2	340	9	0.1	10.50	1224	585	2305		RP	24.16	1.129	0.268	-4.29	0.18	0.15	2.44	39.33	1.655	4.68	24.13	0.094	24.197	0.253	-4.28	0.054	-4.300	0.145	2.57	24.14	0.067	24.197	0.179	-4.28	0.038	-4.300	0.102	2.57	6120	43	55.9	3	2.86	32	8.7	3	4.74	19	0.034	15	0.023	1	X	X	A	19.8	958.1	49.8	447.6	GSEG	Inital Testphase	cover of tension spring is missing	0061_20190621_1159_GPCL3_26088_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
66	GPCL3	26088	GSEG	TRUE	7/10/2019	11:42:14	14:08:49	146.6	55.6875	2968.6500	2	340	169	15.9	10.51	1224	590	2295		RP	19.02	6.708	0.340	-2.51	0.53	0.20	3.07	52.90	9.856	44.15	19.04	0.065	19.152	0.193	-2.54	0.038	-2.583	0.111	7.67	19.04	0.046	19.120	0.137	-2.55	0.027	-2.567	0.078	7.68	114920	1372	57.6	3	3.77	1833	9.1	3	4.99	609	0.030	15	0.019	1	X	X	A	19.7	958.6	49.8	403.4	GSEG	Inital TestPhase after Ringexperiment		0066_20190710_1126_GPCL3_26088_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
76	Leica GPCL3	26088	GSEG	TRUE	10/8/2019	9:08:12	11:36:09	148.0	55.6875	2968.6500	2	340	167	15.9	10.62	1224	590	2295		RP	27.56	4.229	0.342	-2.85	0.69	0.20	3.09	53.14	9.873	44.32	27.54	0.045	27.639	0.194	-2.86	0.026	-2.902	0.111	5.28	27.54	0.032	27.636	0.137	-2.86	0.018	-2.901	0.078	5.28	113560	1674	58.2	3	4.05	1401	9.4	3	5.18	392	0.034	15	0.019	3	X	X	A	19.8	957.1	49.7	428.1	GSEG	Initial Test Phase		0076_20191008_0850_Leica GPCL3_26088_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	26088					24.36	±3.315			-3.45	±0.756

45	GPCL3	26102	GSEG	TRUE	6/17/2019	11:56:54	14:15:45	138.8	55.6875	2968.6500	2	340	9	0.1	10.50	1224	465	2565		RP	43.33	0.433	0.434	-3.88	0.12	0.25	3.97	49.26	1.899	5.41	43.33	0.144	43.334	0.421	-3.86	0.083	-3.888	0.242	3.96	43.32	0.105	43.310	0.307	-3.84	0.061	-3.825	0.177	4.10	6120	38	61.0	3	3.25	65	8.8	3	4.22	60	0.022	15	0.015	1	X	X	B	19.9	958.5	49.7	422.1	GSEG	Inital Test Phase	Rod 26102 seems to be quite deformed. Despite placing the support quite far outside of the Bessel-Points The far end are still 0.3 to 0.4mm below the support points and the center is up to 0.7mm ABOVE the support points.	0045_20190617_1141_GPCL3_26102_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
62	GPCL3	26102	GSEG	TRUE	6/21/2019	16:12:48	18:36:45	144.0	55.6875	2968.6500	2	340	9	0.1	10.50	1224	385	2610		RP	44.50	1.095	0.399	-5.54	0.13	0.23	3.67	46.27	1.944	5.49	44.51	0.138	44.508	0.387	-5.54	0.079	-5.548	0.222	3.80	44.51	0.097	44.508	0.274	-5.54	0.056	-5.548	0.157	3.81	6120	8	55.6	3	2.98	43	7.5	3	4.14	35	0.022	15	0.014	0	X	X	B	19.8	957.2	49.7	428.7	GSEG	Inital Testphase	staff housing is quite deformed. Support points far away from the Bessel Point (both moved towards the outside), but the invar is still 0.8mm above the "support hight".	0062_20190621_1557_GPCL3_26102_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
71	Leica GPCL3	26102	GSEG	TRUE	10/4/2019	8:50:34	11:20:14	149.7	55.6875	2968.6500	2	340	167	15.9	10.61	1224	400	2600		RP	48.63	4.411	0.488	-4.76	1.08	0.28	4.45	54.94	8.962	42.61	48.76	0.057	48.683	0.377	-4.88	0.033	-4.778	0.217	6.76	48.73	0.043	48.681	0.322	-4.85	0.025	-4.775	0.185	7.18	113560	478	59.2	3	3.33	891	8.1	3	4.23	680	0.021	15	0.015	2	X	X	A	19.8	952.8	49.4	431.0	GSEG	Initial Test	Lattenkörper stark deformiert	0071_20191004_0834_Leica GPCL3_26102_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
72	Leica GPCL3	26102	GSEG	TRUE	10/4/2019	12:42:38	15:08:16	145.6	55.6875	2968.6500	2	340	167	15.8	10.61	1224	400	2600		RP	49.53	4.441	0.485	-4.17	1.09	0.28	4.41	54.99	8.979	42.76	49.65	0.057	49.582	0.373	-4.29	0.033	-4.183	0.214	6.74	49.65	0.041	49.584	0.274	-4.28	0.024	-4.184	0.157	6.82	113560	477	59.1	3	3.32	893	8.1	3	4.23	687	0.025	15	0.016	6	X	X	A	19.7	951.5	49.9	450.8	GSEG	Initial Test	Lattenkörper stark deformiert, Wiederholung zu 0071 (leichte Schere Hin/Rück)	0072_20191004_1225_Leica GPCL3_26102_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	26102					46.55	±2.691			-4.63	±0.636

46	LD13	15047	GSEG	TRUE	6/17/2019	14:51:11	16:43:28	112.3	52.2500	2965.0000	2	266	9	0.1	10.50	1224	520	2375		RP	53.04	1.828	0.489	-4.42	0.09	0.29	3.79	-6.09	3.735	10.18	53.07	0.178	53.072	0.419	-4.42	0.104	-4.429	0.244	4.14	53.08	0.126	53.072	0.296	-4.43	0.074	-4.429	0.173	4.15	4788	18	65.0	3	5.47	45	8.2	3	4.11	52	0.029	15	0.017	0	C	B	A	19.9	957.6	49.9	424.6	GSEG	Inital Test Phase		0046_20190617_1434_LD13_15047_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
68	Zeiss LD13	15047	GSEG	TRUE	10/3/2019	9:29:07	11:29:22	120.3	52.2500	2965.0000	2	266	167	15.9	10.62	1224	520	2375		RP	49.40	4.844	0.498	-4.43	2.61	0.29	3.80	-2.99	7.594	48.83	49.52	0.060	49.399	0.268	-4.40	0.035	-4.434	0.156	5.98	49.53	0.043	49.400	0.190	-4.40	0.025	-4.434	0.110	5.98	88844	853	61.2	3	6.28	626	7.8	3	3.99	2001	0.025	15	0.021	1	X	B	A	19.9	960.3	49.7	421.2	GSEG	Initial Test Phase		0068_20191003_0913_Zeiss LD13_15047_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
80	Zeiss LD13	15047	GSEG	TRUE	10/8/2019	16:19:14	18:16:46	117.5	52.2500	2965.0000	2	266	167	15.9	10.63	1224	520	2375		RP	47.73	4.842	0.488	-4.16	2.71	0.28	3.73	-3.11	7.638	49.06	47.82	0.061	47.740	0.244	-4.11	0.036	-4.167	0.142	6.08	47.82	0.043	47.740	0.173	-4.11	0.025	-4.167	0.101	6.08	88844	962	65.6	3	6.72	666	8.6	3	4.17	2140	0.029	15	0.021	2	X	B	A	19.8	953.3	49.9	434.0	GSEG	Inital Test Phase		0080_20191008_1604_Zeiss LD13_15047_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
243	Zeiss LD13	15047	GSEG	FALSE	7/1/2020	9:35:05	11:36:31	121.4	52.2500	2965.0000	2	266	173	15.9	10.30	1224	520	2375	0.000	RP	54.52	5.595	0.535	-4.55	2.72	0.31	4.01	-6.02	9.532	54.75	54.52	0.076	54.409	0.165	-4.50	0.044	-4.585	0.096	7.51	54.52	0.054	54.410	0.119	-4.50	0.031	-4.581	0.069	7.52	92036	1352	60.0	3	8.94	21	9.5	3	3.88	5783	0.061	15	0.079	3	X	A	A	19.6	953.0	50.7	424.1	GSEG	Vorzeichentest	Test bevor Latte zur "Nachkontrolle" nach München kommt	0243_20200701_0913_Zeiss LD13_15047_GSEG.mat	Test before the staff goes to Munich for Sign-Verification
245	Zeiss LD13	15047	GSEG	FALSE	7/1/2020	13:17:41	15:18:31	120.8	52.2500	2965.0000	2	266	205	18.9	10.30	1224	520	2375	0.000	RP	54.42	5.672	0.525	-4.22	3.33	0.30	3.94	-5.99	9.099	55.26	54.37	0.072	54.317	0.186	-4.20	0.042	-4.238	0.108	7.76	54.38	0.051	54.314	0.134	-4.20	0.030	-4.248	0.078	7.77	109060	1527	60.4	3	8.71	35	9.5	3	3.89	6232	0.053	15	0.045	8	X	A	A	19.6	951.6	50.7	406.0	GSEG	Vorzeichentest	Test bevor Latte zur "Nachkontrolle" nach München kommt, Wiederholung zu 0243	0245_20200701_1300_Zeiss LD13_15047_GSEG.mat	Calibration before staff goes to Munich for indipendent sign verification @ TUM
248	Zeiss LD13	15047	GSEG	FALSE	7/9/2020	9:12:54	11:06:10	113.3	52.2500	2965.0000	2	266	173	15.9	10.30	1224	520	2375	0.000	RP	56.94	6.923	0.525	-4.38	3.18	0.30	3.94	-6.07	9.498	54.41	56.89	0.080	57.016	0.152	-4.29	0.047	-4.467	0.089	7.95	56.89	0.057	57.022	0.111	-4.29	0.033	-4.457	0.065	7.95	92036	1481	61.4	3	9.17	28	9.9	3	3.97	5756	0.055	15	0.042	0	X	A	A	19.6	959.0	50.8	430.2	GSEG	Vorzeichentest	Testmeasurement after staff returned from Sitgn-Test in Munich on 07.07.2020	0248_20200709_0859_Zeiss LD13_15047_GSEG.mat	Testmeasurement after the staff returned from munich (sign-test)
249	Zeiss LD13	15047	GSEG	FALSE	9/17/2020	13:47:12	15:40:11	113.0	52.2500	2965.0000	2	266	173	16.0	10.31	1224			0.000	RP	53.01	5.215	0.540	-4.37	2.97	0.31	4.05	-6.08	9.499	54.21	53.05	0.075	52.903	0.165	-4.35	0.044	-4.423	0.096	7.45	53.05	0.053	52.911	0.128	-4.35	0.031	-4.409	0.075	7.47	92036	1428	58.4	3	8.81	28	9.4	3	3.90	6005	0.065	15	0.057	0	X	A	A	19.7	960.3	50.2	434.2	GSEG	Funktionstest Niv-Komp		0249_20200917_1336_Zeiss LD13_15047_GSEG.mat	Functionality Test after break
	Mittelwert	15047					52.75	±2.874			-4.33	±0.127

48	Zeiss LD13	15037	GSEG	TRUE	6/18/2019	7:58:17	9:53:37	115.3	52.2500	2965.0000	2	266	9	0.1	10.50	1224	685	2150		RP	41.72	0.887	0.398	-4.43	0.13	0.23	3.01	-10.88	2.752	7.44	41.73	0.136	41.732	0.343	-4.43	0.079	-4.427	0.200	3.10	41.72	0.097	41.733	0.243	-4.43	0.056	-4.427	0.141	3.11	4788	24	72.7	3	6.14	3	7.6	3	3.41	178	0.036	15	0.028	3	C	B	A	19.9	957.0	49.4	433.0	GSEG	Initial Test Phase	Meas. 0047 interrupted due to Windows Update on Meteo Computer	0048_20190618_0742_Zeiss LD13_15037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
69	Zeiss LD13	15037	GSEG	TRUE	10/3/2019	12:18:22	14:17:40	119.3	52.2500	2965.0000	2	266	167	15.9	10.61	1224	685	2150		RP	41.78	3.507	0.391	-4.03	1.15	0.23	2.98	-5.29	6.910	42.61	41.70	0.044	41.684	0.178	-4.20	0.026	-3.993	0.104	4.32	41.70	0.032	41.724	0.152	-4.19	0.019	-4.022	0.089	4.42	88844	770	67.5	3	5.71	908	6.9	3	3.20	3179	0.040	15	0.034	0	C	A	A	19.9	959.3	49.7	434.7	GSEG	Initial Test Phase		0069_20191003_1202_Zeiss LD13_15037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
79	Zeiss LD13	15037	GSEG	TRUE	10/8/2019	13:10:22	15:07:01	116.7	52.2500	2965.0000	2	266	167	15.9	10.63	1224	685	2150		RP	38.69	5.142	0.393	-4.09	1.29	0.23	3.00	-5.33	6.900	42.86	39.18	0.049	38.694	0.191	-4.28	0.029	-4.086	0.111	4.84	39.18	0.035	38.698	0.146	-4.28	0.021	-4.078	0.085	4.88	88844	797	66.8	3	5.65	893	6.8	3	3.17	2994	0.046	15	0.031	1	C	A	A	19.8	954.9	49.8	440.0	GSEG	Inital Test Phase		0079_20191008_1252_Zeiss LD13_15037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	15037					40.87	±1.193			-4.30	±0.097

TUM2019	Leica GPCL3	76037	GSEG					5.00				3.65
TUM2020	Leica GPCL3	76037	GSEG
50	Leica GPCL3	76037 (Ref.)	GSEG	TRUE	6/18/2019	13:24:56	15:49:48	144.9	55.6875	2968.6500	2	340	9	0.1	10.50	1224	530	2455		RP	-2.90	0.641	0.408	5.58	0.07	0.24	3.68	-5.66	2.132	5.90	-2.87	0.137	-2.931	0.395	5.56	0.079	5.625	0.227	3.71	-2.89	0.098	-2.933	0.282	5.58	0.056	5.627	0.162	3.74	6120	37	69.8	3	5.60	47	8.4	3	3.72	154	0.024	15	0.016	12	C	B	B	19.8	955.3	49.6	449.3	GSEG	Initial Test Phase	measurement 0049 aborted for unknown reason (drive rod was 20 cm away from the drive while (on the right side) ... however it got there???)	0050_20190618_1302_Leica GPCL3_76037 (Ref.)_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
54	GPCL3(Ref.)	76037	GSEG	TRUE	6/19/2019	14:17:28	16:42:14	144.8	55.6875	2968.6500	2	340	9	0.1	10.50	1224	525	2460		RP	-1.34	0.622	0.405	5.76	0.09	0.23	3.68	-5.60	2.178	6.00	-1.34	0.137	-1.319	0.392	5.76	0.079	5.767	0.225	3.73	-1.31	0.098	-1.338	0.281	5.75	0.056	5.775	0.161	3.77	6120	26	73.1	3	5.85	41	9.0	3	3.85	156	0.025	15	0.017	9	C	B	B	19.7	952.2	49.7	428.8	GSEG	Inital Test Phase		0054_20190619_1400_GPCL3(Ref.)_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
59	GPCL3	76037(Ref.)	GSEG	TRUE	6/20/2019	17:21:18	19:45:21	144.0	55.6875	2968.6500	2	340	9	0.1	10.50	1224	530	2455		RP	-0.01	0.653	0.407	5.13	0.08	0.23	3.69	-5.71	2.178	5.94	0.00	0.137	-0.016	0.389	5.13	0.079	5.136	0.223	3.74	0.02	0.097	-0.018	0.276	5.12	0.056	5.137	0.158	3.73	6120	30	71.4	3	5.74	51	8.8	3	3.80	171	0.024	15	0.015	1	C	B	B	19.7	955.5	49.6	432.7	GSEG	Initial Test Phase	measurement 0057 aborted due to an interferometererror (code 11 no sample to read) / measurement 0058 finished, but had an index error while converting from labview to matlab (missing edge very likely)	0059_20190620_1704_GPCL3_76037(Ref.)_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
63	GPCL3	76037	GSEG	TRUE	7/9/2019	11:53:18	14:19:51	146.5	55.6875	2968.6500	2	340	169	15.9	10.49	1224	530	2455		RP	-17.42	4.961	0.399	4.27	1.54	0.23	3.62	-1.47	9.029	39.93	-17.47	0.057	-17.448	0.310	4.28	0.033	4.279	0.178	6.78	-17.47	0.041	-17.445	0.221	4.28	0.023	4.278	0.127	6.79	114920	711	69.3	3	5.80	956	8.3	3	3.64	2572	0.024	15	0.015	1	C	B	A	19.6	957.5	49.7	421.8	GSEG	Inital TestPhase after Ringexperiment		0063_20190709_1136_GPCL3_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
67	GPCL3	76037	GSEG	TRUE	7/10/2019	14:44:08	17:11:08	147.0	55.6875	2968.6500	2	340	169	15.9	10.49	1224	530	2455		RP	-13.05	4.924	0.400	4.19	1.59	0.23	3.64	-1.47	9.047	39.68	-13.10	0.057	-13.071	0.311	4.19	0.033	4.194	0.179	6.76	-13.10	0.040	-13.072	0.220	4.19	0.023	4.194	0.126	6.76	114920	725	67.9	3	5.70	952	8.0	3	3.58	2582	0.023	15	0.013	1	C	B	A	19.7	958.0	49.6	403.7	GSEG	Inital TestPhase after Ringexperiment		0067_20190710_1423_GPCL3_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
82	Leica GPCL3	76037	GSEG	TRUE	10/9/2019	9:45:26	12:08:51	143.4	55.6875	2968.6500	2	340	167	15.9	10.62	1224	530	2455		RP	-16.69	5.103	0.401	4.38	1.68	0.23	3.65	-1.58	9.051	39.94	-16.87	0.056	-16.703	0.311	4.35	0.032	4.385	0.179	6.61	-16.87	0.040	-16.703	0.221	4.35	0.023	4.385	0.127	6.62	113560	655	69.5	3	5.89	831	8.3	3	3.62	2521	0.024	15	0.018	1	C	B	A	19.8	954.1	49.9	430.6	GSEG	Inital Test Phase	0081 interrupted by unknown error of motor controller	0082_20191009_0928_Leica GPCL3_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
101	Leica GPCL3	76037	GSEG	TRUE	10/16/2019	12:21:16	14:49:52	148.6	55.6875	2968.6500	2	340	169	15.9	10.50	1224	530	2455		RP	1.11	4.551	0.399	4.33	1.60	0.23	3.63	-1.46	9.088	39.66	1.08	0.054	1.093	0.316	4.33	0.031	4.344	0.182	6.44	1.08	0.039	1.092	0.228	4.34	0.022	4.345	0.131	6.47	114920	642	70.4	3	5.84	917	8.2	3	3.58	2787	0.023	15	0.015	1	C	B	A	19.7	956.6	49.7	431.7	GSEG	Initial Test Phase		0101_20191016_1206_Leica GPCL3_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
191	Leica GPCL3	76037	GSEG	TRUE	2/24/2020	9:54:40	12:22:53	148.2	55.6875	2968.6500	2	340	463	19.0	10.28	1224	530	2455	0.000	RP	2.26	4.616	0.399	4.18	1.53	0.23	3.61	-1.20	8.901	44.01	2.27	0.030	2.239	0.297	4.20	0.018	4.211	0.170	5.93	2.26	0.022	2.238	0.217	4.20	0.013	4.212	0.124	5.98	314840	1925	64.2	3	5.75	1525	8.3	3	3.50	12638	0.074	15	0.147	2	C	B	A	20.1	962.1	49.4	426.7	GSEG	Nedo Test	Messung für "Nedo Slides" mit erweitertem Codebereich und verdichteten Traces (19mm mit 0.05mm Traceabstand).	0191_20200224_0941_Leica GPCL3_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
192	Leica GPCL3	76037	GSEG	TRUE	2/24/2020	13:22:32	15:51:01	148.5	55.6875	2968.6500	2	340	463	19.0	10.28	1224	530	2455	0.000	RP	11.96	5.028	0.406	4.58	1.52	0.24	3.69	-1.18	8.928	43.92	12.06	0.031	11.983	0.304	4.53	0.018	4.567	0.175	5.97	12.06	0.022	11.983	0.233	4.52	0.013	4.567	0.134	6.08	314840	1790	64.1	3	5.75	1619	8.3	3	3.49	12657	0.096	15	0.208	0	C	C	B	20.0	958.0	49.4	431.0	GSEG	Nedo Test	Messung für "Nedo Slides" mit erweitertem Codebereich und verdichteten Traces (19mm mit 0.05mm Traceabstand). Wiederholung zu 0191, diesmal Kamera-kalibrierung auf "schlechter Kante" (660.150), bei 0191 war die Kalibrierung auf einer "guten Kante" (662.175). Latte unverändert.	0192_20200224_1306_Leica GPCL3_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
193	Leica GPCL3	76037	GSEG	TRUE	2/25/2020	9:46:39	12:15:00	148.3	55.6875	2968.6500	2	340	463	19.0	10.28	1224	530	2455	0.000	RP	2.84	4.602	0.394	4.28	1.53	0.23	3.56	-1.19	8.906	43.87	2.84	0.030	2.838	0.289	4.30	0.018	4.286	0.166	5.90	2.84	0.021	2.819	0.212	4.29	0.012	4.308	0.121	5.93	314840	1987	64.2	3	5.75	1581	8.3	3	3.50	12593	0.069	15	0.086	6	C	B	A	20.0	948.4	49.5	421.5	GSEG	Nedo Tests	Test der Wiederholbarkeit zu 0191. Lattenausrichtung kontrolliert, nicht verändert. Position Nullpunktsmarke "verbessert". Kalibrierung wie bei 0191 an "guter" Kante 662.175mm.	0193_20200225_0933_Leica GPCL3_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
194	Leica GPCL3	76037	GSEG	TRUE	2/25/2020	13:15:17	15:42:47	147.5	55.6875	2968.6500	2	340	463	19.0	10.28	1224	530	2455	0.000	RP	9.97	5.021	0.397	4.49	1.52	0.23	3.59	-1.21	8.925	43.82	10.07	0.030	9.996	0.292	4.43	0.018	4.477	0.168	5.89	10.07	0.021	9.994	0.208	4.43	0.012	4.477	0.120	5.91	314840	1959	64.1	3	5.75	1586	8.3	3	3.49	12574	0.080	15	0.101	3	C	C	B	20.0	946.7	49.5	423.1	GSEG	Nedo Tests	Test der Wiederholbarkeit zu 0192. Lattenausrichtung kontrolliert, nicht verändert. Kalibrierung wie bei 0192 an "schlechter" Kante 660.150mm.	0194_20200225_1303_Leica GPCL3_76037_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
241	Leica GPCL3	76037	GSEG	FALSE	6/30/2020	11:12:18	13:44:22	152.1	55.6875	2968.6500	2	340	173	15.9	10.30	1224	530	2455	0.000	RP	0.34	5.055	0.370	4.02	1.54	0.22	3.35	-0.17	9.316	40.60	0.31	0.051	0.356	0.275	3.99	0.030	4.013	0.158	6.13	0.32	0.037	0.362	0.213	3.98	0.022	4.011	0.122	6.25	117640	860	69.8	3	6.13	762	9.2	3	3.70	5221	0.069	15	0.127	1	C	B	A	19.8	955.2	50.6	405.9	GSEG	Vorzeichentest	Test bevor Latte zur "Nachkontrolle" nach München kommt	0241_20200630_1058_Leica GPCL3_76037_GSEG.mat	Test wit "GSEG Reference Staff" bevore it is sent to Munich for externach "Sign Validation"
242	Leica GPCL3	76037	GSEG	FALSE	6/30/2020	14:46:00	17:17:04	151.1	55.6875	2968.6500	2	340	173	15.9	10.30	1224	530	2455	0.000	RP	-1.39	5.031	0.384	4.33	1.54	0.22	3.47	-0.10	9.330	40.80	-1.41	0.052	-1.397	0.292	4.29	0.030	4.344	0.167	6.19	-1.41	0.037	-1.395	0.210	4.29	0.021	4.345	0.120	6.21	117640	788	65.8	3	5.81	673	8.6	3	3.53	5797	0.066	15	0.110	3	C	B	A	19.7	953.9	50.7	405.9	GSEG	Vorzeichentest	Test bevor Latte zur "Nachkontrolle" nach München kommt, Wiederhlong zu 0241 wegen "Temperaturschere"	0242_20200630_1434_Leica GPCL3_76037_GSEG.mat	Test with GSEG-Reference Staff before it is sent to Munich for sign-verfikation.
247	Leica GPCL3	76037	GSEG	FALSE	7/8/2020	10:21:10	12:38:45	137.6	55.6875	2968.6500	2	340	173	15.9	10.30	1224	530	2455	0.000	RP	-1.03	4.960	0.381	4.13	1.74	0.22	3.45	-0.17	9.310	40.53	-1.01	0.052	-1.044	0.283	4.10	0.030	4.153	0.162	6.17	-1.01	0.037	-1.044	0.203	4.10	0.021	4.153	0.116	6.18	117640	809	66.4	3	5.83	635	8.7	3	3.57	5651	0.067	15	0.058	0	C	B	A	19.8	958.9	50.4	424.0	GSEG	Vorzeichentest	Measurement after Staff returned from Test at TU München (0246 abgebrochen wegen Motorproblem)	0247_20200708_1008_Leica GPCL3_76037_GSEG.mat	Recalibration of staff after it returned from the test in Munich on 0.07.2020.
264	Leica GPCL3	76037	GSEG	FALSE	11/12/2020	9:59:22	12:22:43	143.4	55.6875	2968.6500	2	340	173	16.0	10.31	1224	530	2455	0.000	RP	-0.99	4.951	0.381	4.66	1.56	0.22	3.44	-0.27	9.333	40.79	-1.00	0.052	-0.997	0.288	4.64	0.030	4.675	0.165	6.15	-1.00	0.037	-0.991	0.211	4.63	0.021	4.672	0.121	6.20	117640	783	69.6	3	6.11	725	9.2	3	3.67	5445	0.071	15	0.086	2	C	B	A	20.0	962.7	49.3	449.0	GSEG	Control Measurement		0264_20201112_0943_Leica GPCL3_76037_GSEG.mat	Test with GSEG Reference Staff to check proper operation of the comparator
287	Leica GPCL3	76037	GSEG	FALSE	4/16/2021	14:28:28	16:54:04	145.6	55.6875	2968.6500	2	340	173	15.9	10.30	1224	530	2455	0.000	RP	5.52	4.947	0.382	4.67	1.82	0.22	3.45	-0.23	9.351	40.47	5.52	0.051	5.523	0.285	4.67	0.030	4.682	0.164	6.12	5.53	0.037	5.498	0.219	4.67	0.022	4.707	0.125	6.22	117640	851	66.1	3	5.87	732	8.8	3	3.57	6184	0.097	15	0.200	1	C	B	A	20.1	959.2	49.2	420.9	GSEG	Testlauf	Testlauf mit Referenzlatte nach "Corona-Pause"	0287_20210416_1416_Leica GPCL3_76037_GSEG.mat	Testlauf mit Referenzlatte nach "Corona-Pause"
304	Leica GPCL3	76037	GSEG	FALSE	2/24/2022	13:20:35	15:46:47	146.2	55.6875	2968.6500	2	340	173	15.9	10.30	1224	530	2455	0.000	RP	-1.29	4.618	0.375	4.78	1.71	0.22	3.40	-0.40	9.201	39.32	-1.28	0.049	-1.271	0.290	4.77	0.029	4.779	0.166	5.90	-1.28	0.035	-1.295	0.207	4.77	0.020	4.803	0.119	5.91	117640	739	69.8	3	5.65	754	8.6	3	3.53	5275	0.130	15	0.311	0	C	B	A	19.9	956.9	48.4	460.3	GSEG	Funktionstest	First comparator run after several months with refenece rod to see if compartor is still functioning properly. Repetition to 303 due to one "stuck" edge	0304_20220224_1306_Leica GPCL3_76037_GSEG.mat	none
308	Leica GPCL3	76037	GSEG	FALSE	3/8/2022	10:07:32	12:36:02	148.5	55.6875	2968.6500	2	340	173	16.0	10.32	1224	530	2455	0.000	RP	-2.13	4.871	0.391	4.43	1.67	0.23	3.54	-0.44	9.302	40.57	-2.13	0.052	-2.120	0.298	4.42	0.030	4.438	0.171	6.16	-2.13	0.037	-2.144	0.214	4.42	0.021	4.461	0.123	6.18	117640	740	67.8	3	5.95	680	9.1	3	3.66	5364	0.139	15	0.305	2	C	B	A	20.0	959.3	47.9	442.6	GSEG	Comparatortest		0308_20220308_0954_Leica GPCL3_76037_GSEG.mat	none
332	Leica GPCL3	76037	GSEG	FALSE	3/20/2023	14:58:27	17:25:23	146.9	55.6875	2968.6500	2	340	173	15.9	10.28	1224	530	2455	0.000	RoPr	-1.56	4.654	0.392	3.92	1.61	0.23	3.53	-1.59	9.195	40.13	-1.58	0.050	-1.566	0.300	3.98	0.029	3.945	0.172	5.97	-1.58	0.037	-1.621	0.259	3.98	0.022	4.000	0.148	6.27	117640	792	67.4	3	5.77	842	8.5	3	3.53	5217	0.069	15	0.089	4	C	B	A	19.9	959.9	49.9	451.8	GSEG	internal		0332_20230320_1446_Leica GPCL3_76037_GSEG.mat	none
352	Leica GPCL3	76037	GSEG	FALSE	5/11/2023	10:57:41	13:26:30	148.8	55.6875	2968.6500	2	340	173	15.9	10.27	1224	530	2455	0.000	RoPR	-1.68	4.842	0.383	4.51	1.63	0.22	3.46	-2.14	9.048	38.93	-1.70	0.051	-1.636	0.297	4.51	0.030	4.477	0.170	6.10	-1.71	0.036	-1.636	0.210	4.51	0.021	4.477	0.120	6.11	117640	756	63.4	3	5.02	1231	7.0	3	3.17	5047	0.172	15	0.509	0	C	B	A	19.7	954.3	49.9	466.4	GSEG	Internal	Interferometer geringfügig nachjustiert	0352_20230511_1045_Leica GPCL3_76037_GSEG.mat	none
357	Leica GPCL3	76037	GSEG	FALSE	2/13/2024	13:29:49	16:06:36	156.8	55.6875	2968.6500	2	340	173	15.9	10.27	1224	530	2455	0.000	RoPr	0.11	4.873	0.395	4.80	1.69	0.23	3.56	-2.22	9.053	39.11	0.11	0.052	0.139	0.304	4.81	0.030	4.795	0.175	6.19	0.10	0.037	0.124	0.220	4.81	0.021	4.804	0.126	6.20	117640	744	63.7	3	5.08	1443	7.1	3	3.22	5312	0.619	15	1.203	0	C	B	A	20.1	959.7	49.3	473.8	GSEG	StartUp Test	Initial Test after quite some time not in use!	0357_20240213_1317_Leica GPCL3_76037_GSEG.mat	none
360	Leica GPCL3	76037	GSEG	FALSE	2/21/2024	11:45:46	15:13:10	207.4	55.6875	2968.6500	2	340	173	15.9	10.27	1224	530	2455	0.000	RoPr	4.79	4.823	0.381	4.82	1.67	0.22	3.43	-2.33	9.061	38.71	4.76	0.051	4.821	0.291	4.83	0.030	4.792	0.166	6.08	4.77	0.037	4.832	0.228	4.82	0.022	4.786	0.131	6.23	117640	762	63.9	3	5.08	1421	7.1	3	3.23	5598	0.052	15	0.047	1	C	B	A	20.2	960.2	49.6	481.4	GSEG	none	First Test after replacing motor and bearings	0360_20240221_1133_Leica GPCL3_760377_GSEG.mat	none
361	Leica GPCL3	76037	GSEG	FALSE	2/22/2024	10:31:34	13:59:22	207.8	55.6875	2968.6500	2	340	173	15.9	10.27	1224	530	2455	0.000	RoPr	2.45	4.810	0.392	4.49	1.67	0.23	3.55	-2.31	9.054	38.85	2.42	0.051	2.489	0.308	4.50	0.030	4.465	0.176	6.14	2.43	0.038	2.489	0.247	4.49	0.022	4.465	0.142	6.34	117640	720	63.9	3	5.06	1272	7.1	3	3.20	4621	0.051	15	0.043	0	C	B	A	20.1	943.3	49.6	429.9	GSEG	none	2nd Test after replacing motor and bearings	0361_20240222_1020_Leica GPCL3_760377_GSEG.mat	none
376	Leica GPCL3	76037	GSEG	FALSE	2/26/2025	11:26:33	15:06:54	220.4	55.6875	2968.6500	2	340	171	15.9	10.37	1224	530	2455	0.000	RoPr	-1.09	4.783	0.367	5.00	1.62	0.21	3.33	-1.73	8.970	37.89	-1.04	0.050	-1.064	0.285	5.03	0.029	4.996	0.164	5.90	-1.04	0.035	-1.086	0.207	5.03	0.020	5.019	0.119	5.92	116280	683	68.8	3	5.16	978	7.6	3	3.33	4271	0.077	15	0.059	0	C	B	A	20.2	960.2	49.4	459.3	GSEG	Testrun		0376_20250226_1111_Leica GPCL3_76037_GSEG.mat	none
377	Leica GPCL3	76037	GSEG	FALSE	2/26/2025	16:37:15	20:17:31	220.3	55.6875	2968.6500	2	340	171	15.9	10.37	1224	530	2455	0.000	RoPr	-3.68	4.774	0.378	4.80	1.64	0.22	3.43	-1.92	9.057	38.79	-3.65	0.050	-3.669	0.295	4.84	0.029	4.808	0.169	5.94	-3.65	0.035	-3.647	0.208	4.84	0.021	4.785	0.120	5.95	116280	663	69.6	3	5.32	881	7.8	3	3.38	4363	0.086	15	0.161	1	C	B	A	20.2	960.1	49.4	458.6	GSEG	Testrun	Repetition to 0376	0377_20250226_1622_Leica GPCL3_76037_GSEG.mat	none
378	Leica GPCL3	76037	GSEG	FALSE	2/27/2025	9:18:18	12:58:21	220.1	55.6875	2968.6500	2	340	171	15.9	10.36	1224	530	2455	0.000	RoPr	3.88	4.993	0.398	5.12	1.61	0.23	3.58	-1.97	9.251	40.68	3.91	0.053	3.884	0.301	5.15	0.030	5.123	0.172	6.19	3.91	0.037	3.885	0.217	5.15	0.022	5.123	0.125	6.22	116280	717	70.6	3	6.08	771	9.3	3	3.69	4679	0.076	15	0.077	12	C	B	A	20.1	959.0	49.9	446.0	GSEG	Testrun	Repetition to 0377 (alb Illumination changed (Center: On, Front: On, Rear: OFF)	0378_20250227_0903_Leica GPCL3_76037_GSEG.mat	none
379	Leica GPCL3	76037	GSEG	FALSE	2/27/2025	15:16:55	18:13:10	176.3	55.6875	2968.6500	2	340	173	16.0	10.31	1224	530	2455	0.000	RoPr	-0.21	5.204	0.381	4.78	1.63	0.22	3.44	-1.85	9.313	40.69	-0.27	0.053	-0.179	0.286	4.74	0.031	4.754	0.164	6.27	-0.27	0.037	-0.178	0.203	4.74	0.022	4.754	0.117	6.28	117640	832	66.9	3	5.74	855	8.6	3	3.53	6082	0.065	15	0.060	1	C	B	B	20.2	958.6	49.9	463.7	GSEG	TestRun	Follow Up with slightly raised rod to go back to about 10.27µm/Pix Cameraresolution	0379_20250227_1504_Leica GPCL3_76037_GSEG.mat	none
380	Leica GPCL3	76037	GSEG	FALSE	2/28/2025	10:25:06	13:20:46	175.7	55.6875	2968.6500	2	340	173	15.9	10.26	1224	530	2455	0.000	RoPr	-4.21	5.513	0.383	4.53	1.60	0.22	3.46	-1.87	9.307	40.52	-4.29	0.054	-4.212	0.291	4.52	0.031	4.547	0.167	6.39	-4.29	0.038	-4.214	0.207	4.53	0.022	4.549	0.119	6.39	117640	790	66.6	3	5.70	892	8.5	3	3.55	5973	0.066	15	0.077	4	C	B	A	20.1	961.8	50.0	461.5	GSEG	TestRun	Follow Up with slightly raised rod to go back to about 10.27µm/Pix Cameraresolution (one more step with 0.5mm)	0380_20250228_1011_Leica GPCL3_76037_GSEG.mat	none
	Mittelwert	76037					-0.91	±6.359			4.60	±0.427

89	Leica GPCL0.5 mini	3305	GSEG	TRUE	10/10/2019	15:31:31	15:53:31	22.0	45.5625	426.2625	2	54	163	15.9	10.88	1224	20	480		RP	304.84	3.285	0.567	120.38	14.07	2.25	1.66	11.17	4.806	30.76	304.92	0.086	305.202	0.356	120.32	0.345	119.098	1.429	3.27	304.93	0.061	305.202	0.253	120.26	0.244	119.098	1.014	3.28	17604	265	60.6	3	2.65	738	8.2	3	4.30	113	0.027	15	0.025	0	X	C	C	19.8	960.1	49.9	438.5	GSEG	Inital Test Phase		0089_20191010_1529_Leica GPCL0.5 mini_3305_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
90	Leica GPCL0.5 mini	3305	GSEG	TRUE	10/10/2019	16:28:03	17:34:40	66.6	45.5625	426.2625	6	54	163	15.9	10.87	1224	20	480		RP	304.71	3.231	0.566	120.88	13.90	2.24	1.66	11.17	4.823	30.91	304.72	0.090	305.002	0.346	120.99	0.358	119.850	1.385	3.39	304.72	0.037	304.359	0.120	120.98	0.146	122.035	0.472	3.39	52812	784	60.5	3	2.64	2212	8.1	3	4.31	329	0.025	15	0.015	0	X	C	C	19.8	960.3	49.8	433.7	GSEG	Inital Test Phase	test with multiple iterations	0090_20191010_1611_Leica GPCL0.5 mini_3305_GSEG.mat	Data reprocessed on 03.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	3305					304.82	±0.106			120.62	±0.360

94	Trimble LD13	62398	GSEG	TRUE	10/14/2019	13:36:49	15:36:07	119.3	52.2500	2965.0000	2	266	169	15.9	10.53	1224	550	2450		RP	3.53	6.426	0.564	-2.61	1.50	0.33	4.26	-12.91	6.682	39.63	4.00	0.059	3.510	0.250	-2.55	0.035	-2.580	0.145	5.79	4.01	0.042	3.512	0.181	-2.55	0.024	-2.581	0.105	5.80	89908	691	70.6	3	7.65	256	8.5	3	3.85	3822	0.028	15	0.019	5	X	B	A	19.7	955.2	50.1	462.3	GSEG	Initial Test Phase	0093 interrupted by interferometererror 11	0094_20191014_1319_Trimble LD13_62398_GSEG.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
97	Trimble LD13	62398	GSEG	TRUE	10/15/2019	13:50:08	15:49:56	119.8	52.2500	2965.0000	2	266	169	15.9	10.49	1224	550	2450		RP	1.93	4.425	0.554	-2.66	1.47	0.32	4.21	-12.91	6.657	39.24	2.03	0.058	1.880	0.246	-2.63	0.034	-2.649	0.143	5.76	2.03	0.041	1.880	0.174	-2.63	0.024	-2.649	0.101	5.76	89908	701	71.3	3	7.68	280	8.6	3	3.86	3832	0.026	15	0.018	0	X	B	A	19.7	949.2	49.9	479.3	GSEG	Initial Test Phase		0097_20191015_1331_Trimble LD13_62398_GSEG.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	62398					3.02	±0.988			-2.59	±0.039

95	Trimble LD13	62983	GSEG	TRUE	10/15/2019	8:37:05	10:40:24	123.3	52.2500	2965.0000	2	266	169	15.9	10.53	1224	400	2600		RP	13.89	5.521	0.588	0.66	3.08	0.34	4.46	-22.22	7.098	40.87	13.88	0.063	13.890	0.279	0.72	0.037	0.641	0.162	6.26	13.88	0.045	13.896	0.213	0.72	0.027	0.638	0.124	6.34	89908	782	74.3	3	9.07	71	7.7	3	3.71	3833	0.024	15	0.017	1	X	C	A	19.7	947.7	50.0	472.6	GSEG	Intial Test Phase	Staff with broad foot-plate	0095_20191015_0819_Trimble LD13_62983_GSEG.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
100	Trimble LD13	62983	GSEG	TRUE	10/16/2019	9:26:41	11:27:20	120.7	52.2500	2965.0000	2	266	169	15.8	10.48	1224	400	2600		RP	14.72	5.637	0.576	0.73	3.08	0.34	4.40	-21.67	6.996	40.72	14.69	0.062	14.697	0.284	0.76	0.036	0.734	0.165	6.17	14.69	0.044	14.699	0.202	0.76	0.026	0.734	0.118	6.17	89908	519	78.1	3	8.77	126	8.1	3	3.83	3091	0.024	15	0.014	3	X	C	A	19.7	957.8	49.7	427.1	GSEG	Initial Test Phase		0100_20191016_0910_Trimble LD13_62983_GSEG.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	62983					14.28	±0.401			0.74	±0.021

102	Leica GPCL3	26087	GSEG	TRUE	10/21/2019	9:35:28	11:59:56	144.5	55.6875	2968.6500	2	340	169	15.9	10.51	1224	560	2445		RP	50.94	4.911	0.334	-4.76	0.74	0.19	3.05	45.00	11.694	48.68	50.92	0.047	50.915	0.182	-4.75	0.027	-4.742	0.104	5.59	50.92	0.034	50.951	0.141	-4.75	0.019	-4.758	0.081	5.64	114920	711	62.2	3	3.94	793	7.2	3	4.27	627	0.031	15	0.019	1	X	C	A	19.8	958.5	50.1	460.3	GSEG	Initial Test Phase		0102_20191021_0920_Leica GPCL3_26087_GSEG.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.

105	NEDO GPCL 3	070550 (29)	swisstopo	TRUE	11/11/2019	11:19:54	13:47:59	148.1	55.6875	2968.6500	2	340	169	15.8	10.48	1224	660	2650		RP	-55.29	5.782	0.475	4.14	2.48	0.27	4.29	-29.59	8.088	37.33	-55.29	0.054	-55.314	0.235	4.20	0.031	4.169	0.135	6.31	-55.29	0.038	-55.314	0.174	4.19	0.022	4.169	0.100	6.36	114920	471	75.1	3	7.73	690	7.9	3	5.29	3813	0.035	15	0.023	0	X	C	A	19.8	949.3	49.6	437.8	swisstopo	to be added	oberer Stützpunkt liegt sehr weit aussen. 104 wegen Positionierungsfehler unterbrochen, Latte nicht neu justiert)	0105_20191111_1102_NEDO GPCL 3_070550 (29)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
144	NEDO GPCL3	070550 (29)	swisstopo	TRUE	11/22/2019	9:15:16	11:43:44	148.5	55.6875	2968.6500	2	340	173	15.9	10.24	1224	705	2751		RP	-59.60	6.651	0.425	4.53	2.44	0.24	3.81	-29.62	8.174	37.29	-59.72	0.055	-59.524	0.132	4.64	0.031	4.489	0.076	6.45	-59.71	0.039	-59.529	0.102	4.63	0.022	4.492	0.058	6.48	117640	800	68.0	3	7.18	606	7.1	3	4.98	3876	0.031	15	0.024	4	X	C	A	20.0	948.4	49.6	441.8	swisstopo	5703005198	obere Unterstützung sehr weit vom nominellen Auflagepunkt entfernt, Latte hat aber sehr geringen Durchhang ...	0144_20191122_0856_NEDO GPCL3_070550 (29)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
313	Leica GPCL3	070550 (swisstopo 29)	swisstopo	FALSE	3/10/2022	9:33:15	12:03:17	150.0	55.6875	2968.6500	2	340	171	15.8	10.34	1224	680	2700	0.000	RP	-58.51	6.349	0.452	5.39	2.32	0.26	4.07	-32.60	8.874	42.62	-58.51	0.057	-58.509	0.134	5.40	0.032	5.393	0.077	6.70	-58.52	0.040	-58.514	0.106	5.40	0.023	5.395	0.061	6.73	116280	779	74.8	3	8.60	427	9.1	3	5.58	4262	0.143	15	0.259	4	X	C	A	20.1	963.0	47.6	481.7	swisstopo	5703006202		0313_20220310_0921_Leica GPCL3_070550 (swisstopo 29)_swisstopo.mat	none
314	Leica GPCL3	070550 (swisstopo 29)	swisstopo	FALSE	3/10/2022	13:50:26	16:17:45	147.3	55.6875	2968.6500	2	340	171	15.8	10.34	1224	680	2700	0.000	RP	-56.14	6.359	0.449	5.43	2.23	0.26	4.05	-32.56	8.875	42.55	-56.14	0.056	-56.151	0.130	5.44	0.032	5.448	0.075	6.68	-56.14	0.040	-56.150	0.095	5.44	0.023	5.447	0.055	6.69	116280	739	73.9	3	8.48	403	9.0	3	5.54	4275	0.126	15	0.226	1	X	C	A	20.1	960.4	47.3	451.3	swisstopo	5703006202	Repetition to 313. Latte ordentlich geschüttelt und neu ausgerichtet.	0314_20220310_1339_Leica GPCL3_070550 (swisstopo 29)_swisstopo.mat	none
348	Leica GPCL3	070550 (swisstopo 29)	swisstopo	FALSE	3/28/2023	13:32:35	15:57:46	145.2	55.6875	2968.6500	2	340	173	15.9	10.29	1224	680	2700	0.000	RoPr	-88.81	6.129	0.423	5.29	2.21	0.24	3.79	-31.89	8.259	37.30	-88.88	0.054	-88.762	0.114	5.37	0.031	5.258	0.065	6.38	-88.88	0.038	-88.778	0.084	5.37	0.022	5.264	0.048	6.39	117640	767	65.9	3	6.42	832	6.1	3	4.54	4192	0.098	15	0.079	9	X	C	A	20.2	964.8	49.4	486.7	swisstopo	5703006666		0348_20230328_1319_Leica GPCL3_070550 (swisstopo 29)_swisstopo.mat	none
350	Leica GPCL3	070550 (swisstopo 29)	swisstopo	FALSE	3/29/2023	9:50:09	12:17:36	147.4	55.6875	2968.6500	2	340	173	15.9	10.28	1224	680	2700	0.000	RoPr	-84.37	6.187	0.420	5.50	2.24	0.24	3.78	-31.87	8.250	37.28	-84.45	0.054	-84.337	0.120	5.57	0.031	5.463	0.069	6.41	-84.44	0.038	-84.335	0.088	5.57	0.022	5.462	0.050	6.42	117640	787	65.9	3	6.41	871	6.1	3	4.55	4270	0.103	15	0.099	3	X	C	A	20.2	960.5	49.3	444.5	swisstopo	5703006666		0350_20230329_0938_Leica GPCL3_070550 (swisstopo 29)_swisstopo.mat	none
366	Leica GPCL3	70550 (swisstopo 29)	swisstopo	FALSE	2/27/2024	14:45:29	17:40:26	175.0	55.6875	2968.6500	2	340	173	15.9	10.27	1224	680	2700	0.000	RoPr	-75.68	6.483	0.427	5.87	2.30	0.24	3.85	-32.05	8.276	37.98	-75.67	0.053	-75.657	0.128	5.89	0.030	5.856	0.074	6.27	-75.66	0.037	-75.657	0.092	5.89	0.021	5.856	0.053	6.28	117640	660	73.7	3	7.24	881	7.1	3	5.00	4215	0.085	15	0.051	0	X	C	A	20.1	950.3	49.3	456.0	swisstopo	5703006666 (2024)	small dots from red spraypaint on the lower ca. 15cm of the code. It was not possible to remove them completely wthout the use of any solvents except water and soap.	0366_20240227_1433_Leica GPCL3_70550 (swisstopo 29)_swisstopo.mat	small dots from red spraypaint on the lower ca. 15cm of the code. It was not possible to remove them completely wthout the use of any solvents except water and soap.
367	Leica GPCL3	70550 (swisstopo 29)	swisstopo	FALSE	2/28/2024	10:02:53	12:57:15	174.4	55.6875	2968.6500	2	340	173	15.9	10.27	1224	680	2700	0.000	RoPr	-83.25	6.431	0.423	5.82	2.21	0.24	3.81	-32.09	8.294	38.06	-83.21	0.053	-83.228	0.131	5.84	0.030	5.799	0.075	6.24	-83.21	0.037	-83.228	0.092	5.84	0.021	5.799	0.053	6.24	117640	704	69.5	3	6.88	863	6.6	3	4.78	4199	0.086	15	0.054	0	X	C	A	20.1	959.3	49.4	455.5	swisstopo	5703006666 (2024)	small dots from red spraypaint on the lower ca. 15cm of the code. It was not possible to remove them completely wthout the use of any solvents except water and soap. / Repetition of 366 with invar tape on the top moved to the "left" to see how the scale distribution on the staff changes.	0367_20240228_0951_Leica GPCL3_70550 (swisstopo 29)_swisstopo.mat	small dots from red spraypaint on the lower ca. 15cm of the code. It was not possible to remove them completely wthout the use of any solvents except water and soap.↵Repetition of 366 with invar tape on the top moved to the "left" to see how the scale distribution on the staff changes.
444	Leica GPCL3	70550 (swisstopo 29)	Bundesamt für Landestopografie swisstopo	FALSE	5/8/2025	8:52:09	11:47:01	174.9	55.6875	2968.6500	2	340	173	15.9	10.26	1224	680	2700	0.000	RoPr	-81.31	6.239	0.428	5.87	2.21	0.24	3.87	-31.93	8.583	39.92	-81.36	0.053	-81.282	0.116	5.94	0.031	5.861	0.066	6.36	-81.36	0.038	-81.299	0.083	5.95	0.022	5.870	0.047	6.36	117640	717	70.7	3	8.07	391	7.7	3	5.12	4404	0.080	15	0.061	0	X	C	A	19.7	952.6	50.6	452.8	Bundesamt für Landestopografie swisstopo	9950001277	Small code defects, scratches @1260 & 1760mm.	0444_20250508_0841_Leica GPCL3_70550 (swisstopo 29)_Bundesamt für Landestopografie swisstopo.mat	Small code defects, scratches @1260 & 1760mm.
					-70.23	±13.311			5.29	±0.549

108	NEDO GPCL 3	070551 (30)	swisstopo	TRUE	11/11/2019	15:15:12	17:41:24	146.2	55.6875	2968.6500	2	340	171	15.9	10.42	1224	500	2550		RP	-18.49	5.132	0.515	4.25	1.75	0.30	4.65	-29.49	8.108	39.82	-18.61	0.057	-18.489	0.242	4.33	0.033	4.251	0.138	6.68	-18.60	0.040	-18.487	0.181	4.33	0.023	4.249	0.104	6.74	116280	688	74.7	3	7.09	844	8.4	3	5.47	3840	0.029	15	0.023	2	X	C	A	19.8	948.0	49.5	466.1	swisstopo	5703005198	106 unterbrochen wegen Interferometerfehler, Latte nicht neu ausgerichtet, 107 abgebrochen wegen sehr inhomogener Kante für Kalibrierung	0108_20191111_1457_NEDO GPCL 3_070551 (29)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
146	NEDO GPCL3	070551 (30)	swisstopo	TRUE	11/22/2019	13:54:00	16:18:30	144.5	55.6875	2968.6500	2	340	173	15.9	10.24	1224	487	2530		RP	-18.65	4.523	0.468	3.78	1.83	0.27	4.22	-29.57	8.226	41.43	-18.59	0.049	-18.629	0.120	3.94	0.028	3.764	0.069	5.78	-18.59	0.035	-18.609	0.098	3.94	0.020	3.755	0.056	5.83	117640	747	68.9	3	6.79	735	7.7	3	5.23	3945	0.025	15	0.016	0	X	C	A	19.9	945.1	49.6	456.2	swisstopo	5703005198	Messung 0145 unterbrochen wegen Interferometerfehler ... Latte nicht neu justiert, Messung nur neu gestartet.	0146_20191122_1337_NEDO GPCL3_070551 (30)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
315	Leica GPCL3	070551 (swisstopo 30)	swisstopo	FALSE	3/11/2022	8:45:29	11:12:34	147.1	55.6875	2968.6500	2	340	173	15.9	10.29	1224	500	2550	0.000	RP	-15.22	5.007	0.506	5.01	0.73	0.29	4.58	-32.76	9.047	47.45	-15.23	0.055	-15.203	0.110	5.01	0.031	5.015	0.063	6.53	-15.22	0.039	-15.203	0.079	5.01	0.022	5.015	0.045	6.54	117640	903	74.8	3	8.53	419	9.8	3	6.01	4421	0.089	15	0.218	0	X	C	A	20.1	958.2	47.7	462.9	swisstopo	5703006202		0315_20220311_0834_Leica GPCL3_070551 (swisstopo 30)_swisstopo.mat	none
316	Leica GPCL3	070551 (swisstopo 30)	swisstopo	FALSE	3/11/2022	13:22:10	15:50:03	147.9	55.6875	2968.6500	2	340	173	15.9	10.29	1224	500	2550	0.000	RP	-19.48	4.424	0.491	5.11	0.88	0.28	4.44	-32.56	9.068	46.87	-19.38	0.051	-19.479	0.101	5.13	0.029	5.126	0.058	6.13	-19.38	0.036	-19.468	0.075	5.13	0.021	5.123	0.043	6.13	117640	921	74.7	3	8.53	485	9.8	3	6.08	4446	0.094	15	0.185	0	X	C	A	20.2	955.4	47.6	448.6	swisstopo	5703006202	Wiederholung zu 315. Latte ordentlich geschüttelt und neu ausgerichtet.	0316_20220311_1310_Leica GPCL3_070551 (swisstopo 30)_swisstopo.mat	none
346	Leica GPCL3	70551 (swisstopo 30)	swisstopo	FALSE	3/27/2023	14:14:46	16:36:52	142.1	55.6875	2968.6500	2	340	173	15.9	10.28	1224	500	2550	0.000	RoPr	-23.59	5.133	0.475	4.95	0.98	0.27	4.28	-31.89	8.248	41.16	-23.65	0.053	-23.590	0.113	5.01	0.030	4.957	0.065	6.24	-23.65	0.037	-23.536	0.080	5.01	0.021	4.932	0.046	6.25	117640	682	66.2	3	6.14	988	6.7	3	4.84	4290	0.095	15	0.156	1	X	C	A	20.1	960.0	49.4	436.4	swisstopo	5703006666		0346_20230327_1359_Leica GPCL3_70551 (swisstopo 30)_swisstopo.mat	none
347	Leica GPCL3	070551 (swisstopo 30)	swisstopo	FALSE	3/28/2023	10:06:49	12:31:33	144.7	55.6875	2968.6500	2	340	173	15.9	10.28	1224	500	2550	0.000	RoPr	-19.82	4.706	0.473	4.91	0.96	0.27	4.26	-31.75	8.206	41.09	-19.68	0.050	-19.724	0.114	4.94	0.029	4.861	0.065	5.92	-19.69	0.035	-19.708	0.084	4.95	0.020	4.853	0.048	5.93	117640	662	66.3	3	6.14	958	6.8	3	4.86	4429	0.094	15	0.167	0	X	C	A	20.1	966.3	49.5	469.5	swisstopo	5703006666		0347_20230328_0953_Leica GPCL3_070551 (swisstopo 30)_swisstopo.mat	none
363	Leica GPCL3	70551 (swisstopo 30)	swisstopo	FALSE	2/26/2024	9:44:25	12:39:56	175.5	55.6875	2968.6500	2	340	173	15.9	10.29	1224	500	2550	0.000	RoPr	-31.72	4.329	0.477	5.47	1.39	0.27	4.29	-31.44	8.380	40.75	-31.71	0.051	-31.694	0.107	5.57	0.029	5.442	0.061	6.05	-31.71	0.036	-31.693	0.076	5.57	0.021	5.442	0.043	6.05	117640	631	70.6	3	6.44	1043	7.4	3	5.11	4250	0.070	15	0.046	1	X	C	A	20.2	939.2	49.6	440.6	swisstopo	5703006666 (2024)		0363_20240226_0933_Leica GPCL3_70551 (swisstopo 30)_swisstopo.mat	none
365	Leica GPCL3	70551 (swisstopo 30)	swisstopo	FALSE	2/27/2024	9:59:45	12:55:22	175.6	55.6875	2968.6500	2	340	173	15.9	10.30	1224	500	2550	0.000	RoPr	-33.43	4.303	0.476	5.16	1.40	0.27	4.29	-31.43	8.374	40.79	-33.42	0.051	-33.393	0.111	5.26	0.029	5.137	0.064	6.04	-33.42	0.036	-33.331	0.079	5.26	0.021	5.105	0.045	6.05	117640	591	69.8	3	6.38	1025	7.2	3	5.04	4205	0.074	15	0.053	1	X	C	A	20.1	949.3	49.4	451.8	swisstopo	5703006666 (2024)	Wiederholung zu 0363 (Unterschied Nullpunkt?)	0365_20240227_0948_Leica GPCL3_70551 (swisstopo 30)_swisstopo.mat	none
440	Leica GPCL3	70551 (swisstopo 30)	Bundesamt für Landestopografie swisstopo	FALSE	5/5/2025	11:24:57	14:20:24	175.5	55.6875	2968.6500	2	340	173	15.9	10.26	1224	500	2550	0.000	RoPr	-38.09	4.313	0.483	4.56	0.77	0.28	4.33	-31.61	8.447	41.76	-37.95	0.051	-38.093	0.110	4.53	0.029	4.564	0.063	5.98	-37.95	0.036	-38.082	0.080	4.53	0.020	4.560	0.046	5.99	117640	593	59.9	3	5.76	606	6.7	3	4.52	4343	0.074	15	0.052	4	X	C	A	19.7	949.4	50.7	446.2	Bundesamt für Landestopografie swisstopo	9950001277	Small scratch @ 2495mm. / Damaged code edge @ 900 & 2687mm.	0440_20250505_1110_Leica GPCL3_70551 (swisstopo 30)_Bundesamt für Landestopografie swisstopo.mat	Small scratch @ 2495mm. Damaged code edge @ 900 & 2687mm.
442	Leica GPCL3	70551 (swisstopo 30)	Bundesamt für Landestopografie swisstopo	FALSE	5/6/2025	16:06:05	19:00:33	174.5	55.6875	2968.6500	2	340	173	15.9	10.26	1224	500	2550	0.000	RoPr	-35.37	4.369	0.485	5.74	0.72	0.28	4.36	-31.44	8.415	41.80	-35.23	0.051	-35.359	0.098	5.71	0.029	5.725	0.056	6.05	-35.23	0.036	-35.389	0.071	5.71	0.021	5.739	0.041	6.05	117640	591	70.5	3	6.78	622	8.2	3	5.22	4304	0.083	15	0.055	3	X	C	A	19.7	953.8	50.8	442.6	Bundesamt für Landestopografie swisstopo	9950001277	Small scratch @ 2495mm. Damaged code edge @ 900 & 2687mm.	0442_20250506_1554_Leica GPCL3_70551 (swisstopo 30)_Bundesamt für Landestopografie swisstopo.mat	Small scratch @ 2495mm. Damaged code edge @ 900 & 2687mm.
					-25.34	±7.914			4.94	±0.518

109	NEDO GPCL3	070552 (swisstopo 33)	swisstopo	TRUE	11/12/2019	8:32:45	10:58:11	145.4	55.6875	2968.6500	2	340	171	15.9	10.40	1224	550	2645		RP	-18.86	5.206	0.404	3.09	1.72	0.23	3.66	-24.27	7.834	39.66	-18.58	0.051	-18.891	0.237	2.97	0.029	3.123	0.136	5.97	-18.58	0.036	-18.893	0.169	2.97	0.021	3.124	0.097	5.97	116280	684	70.4	3	7.60	404	7.4	3	4.59	2991	0.029	15	0.019	5	X	C	A	19.8	948.9	49.4	432.8	swisstopo	5703005198		0109_20191112_0815_NEDO GPCL3_070552 (swisstopo 33)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
147	NEDO GPCL3	070552 (33)	swisstopo	TRUE	11/25/2019	8:59:12	11:28:43	149.5	55.6875	2968.6500	2	340	173	15.8	10.23	1224	555	2648		RP	-19.74	5.496	0.351	3.40	1.55	0.20	3.18	-24.43	7.995	41.10	-19.41	0.050	-19.760	0.151	3.33	0.029	3.419	0.087	5.89	-19.41	0.035	-19.737	0.107	3.33	0.020	3.407	0.061	5.89	117640	870	67.8	3	7.64	428	7.2	3	4.55	2974	0.032	15	0.026	5	X	C	A	20.0	950.0	49.5	448.1	swisstopo	5703005198		0147_20191125_0844_NEDO GPCL3_070552 (33)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
					-18.99	±0.414			3.15	±0.182

111	Zeiss LD13	10200 (swisstopo 35)	swisstopo	TRUE	11/12/2019	15:17:32	17:16:28	118.9	52.2500	2965.0000	2	266	171	15.9	10.42	1224	750	2700		RP	-23.05	6.726	0.894	-0.57	3.94	0.52	6.80	-6.29	7.734	40.16	-22.59	0.080	-23.110	0.797	-0.94	0.047	-0.497	0.464	7.94	-22.59	0.057	-23.110	0.563	-0.94	0.033	-0.497	0.328	7.95	90972	299	65.6	3	5.00	473	4.6	3	2.33	2896	0.030	15	0.018	0	X	B	B	19.9	945.7	49.5	462.4	swisstopo	5703005198	Gehäuse hat einen sichtbaren Schlag bei ca. 2.02m daher stark deformierter Lattenkörper, Durchbiegung lässt sich durch Veränderung der Lagerposition nicht ausgleichen!	0111_20191112_1458_Zeiss LD13_10200 (swisstopo 35)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
151	Zeiss LD13	010200 (35)	swisstopo	TRUE	11/26/2019	8:19:16	10:16:32	117.3	52.2500	2965.0000	2	266	173	15.9	10.25	1224	677	2701		RP	-38.92	6.795	0.381	2.26	3.58	0.22	2.81	-6.30	7.701	39.39	-38.59	0.053	-39.071	0.207	2.11	0.031	2.327	0.120	5.14	-38.60	0.038	-39.076	0.153	2.11	0.022	2.329	0.089	5.17	92036	1821	64.4	3	5.02	546	4.6	3	2.41	3284	0.030	15	0.018	2	C	B	A	20.0	950.3	49.8	465.0	swisstopo	5703005198	Latte zeigt Knick im Gehäuse bei rund 2m, daher stark deformiert	0151_20191126_0800_Zeiss LD13_010200 (35)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Minor damage at staff housing at approximately 2m, visible only when viewed at an oblique angle
					-30.59	±8.007			0.59	±1.528

112	Zeiss LD13	010154 (swisstopo 38)	swisstopo	TRUE	11/14/2019	10:45:06	12:43:51	118.8	52.2500	2965.0000	2	266	171	15.9	10.40	1224	560	2565		RP	19.39	3.758	0.494	0.25	2.33	0.29	3.78	0.62	5.706	30.99	19.50	0.052	19.029	0.294	0.29	0.030	0.785	0.175	5.21	19.49	0.037	19.041	0.212	0.29	0.022	0.757	0.126	5.22	90972	1700	65.9	3	5.51	417	7.4	3	3.70	975	0.025	15	0.016	1	C	A	A	19.9	942.1	49.9	466.0	swisstopo	5703005198	Sichtbarer Schlag im Gehäuse bei ca. 2.49m, Latte "geknickt" bei ca 1950mm?	0112_20191114_1027_Zeiss LD13_010154 (swisstopo 38)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Visible Damage at approximately 2.5m
143	Zeiss LD13	010154 (swisstopo 38)	swisstopo	TRUE	11/21/2019	16:09:33	17:11:04	61.5	55.0000	2950.0000	2	110	173	15.9	10.27	1224	670	2331		RP	17.93	4.325	0.753	-0.32	2.58	0.44	3.73	0.31	5.634	30.36	18.30	0.090	17.733	0.540	-0.50	0.053	-0.247	0.318	5.88	18.33	0.064	17.733	0.384	-0.51	0.037	-0.247	0.226	5.87	38060	514	60.3	3	4.79	164	6.8	3	3.99	312	0.080	15	0.251	0	C	A	A	19.9	948.7	49.5	485.0	swisstopo	Camera Alignment Test	schneller Durchlauf nach Kamerajustierung	0143_20191121_1554_Zeiss LD13_010154 (swisstopo 38)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
157	Zeiss LD13	010154 (38)	swisstopo	TRUE	11/27/2019	8:21:02	10:18:10	117.1	52.2500	2965.0000	2	266	173	15.8	10.22	1224	562	2565		RP	12.05	3.579	0.424	0.41	2.27	0.25	3.20	0.60	5.694	31.09	12.26	0.048	11.953	0.269	0.33	0.028	0.456	0.157	4.82	12.26	0.034	11.953	0.194	0.33	0.020	0.456	0.113	4.84	92036	1390	61.1	3	5.20	387	6.9	3	3.79	847	0.025	15	0.018	1	C	A	A	20.1	942.8	49.6	466.1	swisstopo	5703005198	Latte mit Schlag im Gehäuse bei ca. 2.45m. Messung 0154 durch Stromausfall unterbrochen, 0155 wegen Kamerafehler unterbrochen, 0156 wegen Motorfehler unterbrochen. Lattenausrichtung überprüft, aber nicht neu justiert	0157_20191127_0806_Zeiss LD13_010154 (38)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Damage on staff housing at approximately 2.5m visible, code-tape still movable but shows abrasion from guide-groove
					16.69	±3.170			0.03	±0.387

113	Zeiss LD13	010188 (swisstopo 37)	swisstopo	TRUE	11/14/2019	14:04:37	16:01:45	117.1	52.2500	2965.0000	2	266	171	15.9	10.41	1224	640	2090		RP	-18.55	5.001	0.640	4.02	2.99	0.38	4.87	-2.00	6.557	34.68	-18.54	0.073	-18.584	0.525	4.04	0.043	4.034	0.305	7.26	-18.54	0.052	-18.582	0.379	4.04	0.030	4.033	0.220	7.31	90972	190	65.1	3	5.55	579	5.3	3	2.48	4418	0.035	15	0.027	8	C	B	A	19.9	939.8	49.5	477.3	swisstopo	5703005198	Latte biegt sich stark durch ... weiches Gehäuse???	0113_20191114_1349_Zeiss LD13_010188 (swisstopo 37)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
152	Zeiss LD13	010188 (37)	swisstopo	TRUE	11/26/2019	11:05:00	13:03:09	118.2	52.2500	2965.0000	2	266	173	15.8	10.24	1224	645	2086		RP	-28.84	4.583	0.426	3.57	3.02	0.25	3.23	-1.79	6.575	34.26	-28.71	0.068	-28.841	0.260	3.56	0.040	3.574	0.152	6.79	-28.71	0.048	-28.841	0.189	3.56	0.028	3.574	0.110	6.81	92036	627	62.7	3	5.49	458	5.2	3	2.43	5078	0.038	15	0.035	0	C	A	A	20.0	949.0	49.7	479.8	swisstopo	5703005198	Latte zeigt überraschend grossen, "symmetrischen" Durchhang ... weiches Gehäuse?	0152_20191126_1048_Zeiss LD13_010188 (37)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Strong signs of wear visible High staff-deflection on support
					-23.62	±5.086			3.80	±0.242

114	Zeiss LD12	15546 (swisstopo 98)	swisstopo	TRUE	11/15/2019	9:24:18	10:42:49	78.5	52.2500	1925.0000	2	175	171	15.9	10.39	1224	530	1475		RP	60.57	4.530	0.403	-4.53	3.28	0.35	2.37	6.81	8.321	43.00	60.68	0.065	60.823	0.300	-4.63	0.057	-4.759	0.262	5.01	60.68	0.046	60.819	0.215	-4.63	0.040	-4.756	0.188	5.03	59850	855	68.8	3	4.62	497	5.0	3	3.60	1621	0.028	15	0.021	1	C	B	A	19.9	940.6	49.7	444.4	swisstopo	5703005198	Optimierung des Durchhanges auf Grund der Thermometerhalterung nicht möglich.	0114_20191115_0908_Zeiss LD12_15546 (swisstopo 98)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
159	Zeiss LD12	15546 (98)	swisstopo	TRUE	11/27/2019	13:27:03	14:43:38	76.6	52.2500	1925.0000	2	175	173	15.8	10.22	1224	535	1485		RP	39.28	4.305	0.308	-2.45	4.01	0.27	1.80	6.95	8.292	42.95	38.71	0.067	39.034	0.119	-1.92	0.059	-2.297	0.103	5.14	38.71	0.047	39.023	0.085	-1.92	0.042	-2.289	0.074	5.14	60550	936	68.5	3	4.45	582	4.9	3	3.36	2314	0.026	15	0.014	0	C	B	A	20.1	940.6	49.6	465.9	swisstopo	5703005198		0159_20191127_1301_Zeiss LD12_15546 (98)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
373	Zeiss LD12	15546 (swisstopo 98)	swisstopo	FALSE	3/1/2024	13:59:15	15:31:05	91.8	52.2500	1935.0000	2	176	173	15.9	10.26	1224	535	1410	0.000	RoPr	45.05	6.324	0.306	-3.52	3.88	0.27	1.81	6.23	8.433	45.32	44.29	0.074	44.849	0.113	-3.27	0.065	-3.423	0.098	5.70	44.29	0.053	44.952	0.105	-3.27	0.046	-3.503	0.092	5.72	60896	1897	66.2	3	5.31	1056	5.0	3	4.39	2022	0.054	15	0.038	1	C	B	A	20.1	947.7	49.6	484.0	swisstopo	5703006666 (2024)	Clearly visible dent @1.0m on the right side of the staff housing. / support can not be placed in exactly the same positions as @ 158 due to auxiliary things on the back of the staff.	0373_20240301_1347_Zeiss LD12_15546 (swisstopo 98)_swisstopo.mat	Clearly visible dent @1.0m on the right side of the staff housing.
375	Zeiss LD12	15546 (swisstopo 98)	swisstopo	FALSE	3/4/2024	9:59:38	11:32:01	92.4	52.2500	1935.0000	2	176	173	15.9	10.26	1224	535	1410	0.000	RoPr	32.71	8.895	0.309	-3.68	3.83	0.27	1.83	6.14	8.439	44.97	30.97	0.072	32.366	0.127	-3.26	0.063	-3.506	0.111	5.53	30.96	0.051	32.495	0.093	-3.25	0.045	-3.604	0.081	5.53	60896	1858	66.2	3	5.20	1107	5.0	3	4.49	1941	0.054	15	0.049	0	C	B	A	20.1	951.2	49.8	436.5	swisstopo	5703006666 (2024)	Clearly visible dent @1.0m on the right side of the staff housing. / support can not be placed in exactly the same positions as @ 158 due to auxiliary things on the back of the staff. / Repetition to 0373	0375_20240304_0949_Zeiss LD12_15546 (swisstopo 98)_swisstopo.mat	Clearly visible dent @1.0m on the right side of the staff housing.
					43.66	±10.906			-3.27	±0.957

115	Zeiss LD12	15544 (swisstopo 97)	swisstopo	TRUE	11/15/2019	11:37:24	12:56:36	79.2	52.2500	1925.0000	2	175	171	15.9	10.39	1224	280	1565		RP	28.16	4.204	0.332	-2.70	3.19	0.29	1.85	5.57	6.517	30.62	27.81	0.060	27.833	0.164	-2.49	0.052	-2.519	0.143	4.44	27.81	0.042	27.917	0.115	-2.50	0.037	-2.579	0.099	4.45	59850	1730	64.8	3	3.85	1995	6.1	3	4.69	2841	0.018	15	0.012	2	C	B	A	20.0	941.1	49.6	455.7	swisstopo	5703005198	Optimierung des Durchhanges auf Grund der Thermometerhalterung bzw. d. untern Dosenlibelle nicht möglich.	0115_20191115_1120_Zeiss LD12_15544 (swisstopo 97)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
158	Zeiss LD12	15544 (97)	swisstopo	TRUE	11/27/2019	11:00:34	12:16:50	76.3	52.2500	1925.0000	2	175	175	16.0	10.19	1224	211	1571		RP	23.51	3.797	0.300	0.23	2.73	0.26	1.71	5.62	6.502	30.48	23.31	0.050	23.284	0.123	0.31	0.043	0.339	0.108	3.85	23.31	0.035	23.285	0.089	0.30	0.031	0.337	0.078	3.85	61250	1034	64.3	3	3.84	1964	6.2	3	4.83	2862	0.018	15	0.012	1	C	A	A	20.1	941.2	49.6	468.1	swisstopo	5703005198		0158_20191127_1045_Zeiss LD12_15544 (97)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
371	Zeiss LD12	15544 (swisstopo 97)	swisstopo	FALSE	3/1/2024	10:30:48	12:02:41	91.9	52.2500	1935.0000	2	176	173	15.9	10.27	1224	210	1630	0.000	RoPr	16.78	3.596	0.293	-0.86	2.19	0.26	1.67	5.55	6.552	30.61	16.55	0.050	16.567	0.118	-0.67	0.043	-0.685	0.103	3.82	16.54	0.035	16.567	0.089	-0.67	0.031	-0.685	0.078	3.83	60896	1220	62.8	3	3.52	2474	5.7	3	4.49	3417	0.032	15	0.023	0	C	A	A	20.1	948.9	49.4	487.8	swisstopo	5703006666 (2024)	support can not be placed in exactly the same positions as @ 158 due to auxiliary things on the back of the staff.	0371_20240301_1017_Zeiss LD12_15544 (swisstopo 97)_swisstopo.mat	none
					22.56	±4.632			-0.95	±1.159

118	Leica GPCL2	26494 (swisstopo 91)	swisstopo	TRUE	11/15/2019	14:57:45	16:22:46	85.0	55.6875	1729.3500	2	190	173	15.9	10.30	1224	125	1601		RP	0.01	4.238	0.745	-5.06	6.04	0.77	5.09	18.47	8.465	38.05	0.05	0.097	0.202	0.692	-5.15	0.100	-5.190	0.715	8.76	0.04	0.069	0.202	0.489	-5.14	0.071	-5.189	0.505	8.77	65740	782	60.6	3	3.47	1071	7.6	3	4.51	420	0.017	15	0.011	0	X	C	B	20.0	941.3	49.6	478.4	swisstopo	5703005198	Codeband klemmt bei ca. 253mm, Lattenkörper stark verbogen und tordiert. Sonderlatte gekürzt auf 1.80m	0118_20191115_1442_Leica GPCL2_26494 (swisstopo 91)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
160	WILD GPCL2	26494 (91)	swisstopo	TRUE	11/27/2019	15:47:56	17:08:17	80.3	55.6875	1729.3500	2	190	173	15.8	10.22	1224	125	1615		RP	17.90	4.739	0.348	-4.26	7.62	0.36	2.38	18.60	8.452	37.76	18.06	0.074	17.750	0.187	-4.76	0.077	-4.104	0.193	6.72	18.05	0.053	17.748	0.131	-4.76	0.054	-4.102	0.135	6.72	65740	769	57.8	3	3.22	1064	7.1	3	4.39	420	0.017	15	0.011	0	X	B	A	20.0	940.6	49.6	469.2	swisstopo	5703005198	Lattenband klemmt bei 250mm, Schraube links unten im Fuss fehlt. Lattenkörper liegt sehr schlecht auf (Verdrehung des Gehäuses)	0160_20191127_1533_WILD GPCL2_26494 (91)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Small indention of the code tape guide at approximately 25cm, code tape not movable anymore top bottom right screw of staff foot is missing
					9.05	±9.006			-4.95	±0.190

120	WILD GPCL2	26493 (swisstopo 92)	swisstopo	TRUE	11/18/2019	15:22:27	16:46:00	83.5	55.6875	1721.2500	2	190	171	16.0	10.44	1224	256	1490		RP	-82.95	3.085	0.508	-5.24	4.21	0.52	3.48	17.73	6.371	33.44	-83.03	0.060	-82.987	0.457	-5.14	0.061	-5.222	0.473	5.35	-83.03	0.042	-82.987	0.322	-5.14	0.043	-5.222	0.334	5.35	64980	434	61.0	3	2.92	1331	7.0	3	4.08	449	0.018	15	0.013	0	C	C	A	20.0	952.1	49.6	449.6	swisstopo	5703005198	0119 wegen Interferometerfehler abgebrochen. Latte nicht neu justiert (wie vorher: Minimierung Durchhang wegen Griffen, Libellen und Halterung für Temperatursensor nicht möglich).	0120_20191118_1504_WILD GPCL2_26493 (swisstopo 92)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
161	WILD GPCL2	26493 (92)	swisstopo	TRUE	11/28/2019	7:38:38	8:58:17	79.6	55.6875	1703.0250	2	189	173	15.8	10.22	1224	262	1492		RP	-90.44	3.589	0.242	-3.32	4.86	0.25	1.64	17.64	6.469	33.43	-90.54	0.057	-90.446	0.127	-2.58	0.059	-3.309	0.132	5.08	-90.53	0.040	-90.434	0.101	-2.60	0.042	-3.320	0.105	5.11	65394	404	59.8	3	2.77	1078	7.1	3	4.20	165	0.018	15	0.014	4	C	B	A	20.1	942.4	49.7	439.7	swisstopo	5703005198	auf 1.80m gekürzte Latte	0161_20191128_0719_WILD GPCL2_26493 (92)_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
					-86.78	±3.751			-3.87	±1.269

121	Zeiss GWLD182	swisstopo 99	swisstopo	TRUE	11/19/2019	10:22:25	11:34:47	72.4	57.7500	1815.0000	2	165	173	16.0	10.31	1224	421	1431		RP	60.42	3.546	1.558	6.47	4.79	1.43	8.97	1.83	7.684	35.50	60.35	0.142	60.141	1.524	6.56	0.130	6.628	1.398	10.79	60.35	0.100	60.141	1.074	6.56	0.092	6.628	0.985	10.79	57090	32	57.8	3	3.02	604	5.6	3	3.65	490	0.019	15	0.013	1	C	B	B	20.0	956.3	49.6	493.1	swisstopo	5703005198	Güteklasse A Banane	0121_20191119_1006_Zeiss GWLD182_swisstopo 99_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
123	Zeiss GWLD182	swisstopo 99	swisstopo	TRUE	11/19/2019	14:03:14	15:17:40	74.4	57.7500	1815.0000	2	165	171	15.9	10.37	1224	655	1149		RP	43.77	4.360	0.968	6.40	5.40	0.88	5.47	1.76	7.704	35.63	43.61	0.106	43.484	0.922	6.57	0.097	6.593	0.844	7.89	43.60	0.075	43.482	0.651	6.58	0.069	6.592	0.595	7.89	56430	460	59.0	3	3.29	424	5.7	3	3.83	409	0.021	15	0.017	9	C	B	B	20.0	955.6	49.5	495.7	swisstopo	5703005198	Güteklasse A Banane, Wiederholung mit stark veränderten Auflagerpunkten (122 wegen Interferometerfehler agbebroche, Latte nicht neu justiert).	0123_20191119_1347_Zeiss GWLD182_swisstopo 99_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
162	Zeiss GWLD182	99	swisstopo	TRUE	11/28/2019	10:33:02	11:45:28	72.4	52.2500	1815.0000	2	166	173	16.0	10.32	1224	423	1432		RP	36.47	3.396	0.509	4.75	4.86	0.47	2.91	1.80	7.736	36.09	36.50	0.091	36.578	0.397	4.73	0.084	4.674	0.365	6.91	36.50	0.065	36.578	0.280	4.73	0.059	4.674	0.258	6.91	57436	329	65.3	3	3.80	529	6.6	3	4.18	353	0.018	15	0.010	6	C	A	A	20.1	943.5	49.3	517.0	swisstopo	5703005198	sehr stark gekrümmt, Auflageroptimierung unöglich	0162_20191128_1017_Zeiss GWLD182_99_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. No obvious defects but staff shows high deflection on support

293	Leica GPCL3	78911 (swisstopo 39)	swisstopo	FALSE	4/26/2021	9:53:45	12:20:15	146.5	55.6875	2968.6500	2	340	173	15.9	10.27	1224	510	2500	0.000	RP	6.39	4.019	0.409	4.78	2.87	0.23	3.67	4.25	3.258	18.69	6.19	0.048	6.308	0.217	4.79	0.028	4.799	0.124	5.69	6.19	0.035	6.311	0.187	4.78	0.020	4.799	0.107	5.87	117640	573	61.8	3	6.64	131	8.4	3	3.39	5038	0.082	15	0.199	0	C	B	A	20.0	954.8	50.0	423.2	swisstopo	Vergleich TUM	ungebrauchte Latte, Verglcieh mit TUM	0293_20210426_0942_Leica GPCL3_78911 (swisstopo 39)_swisstopo.mat	none
317	Leica GPCL3	78911 (swisstopo 39)	swisstopo	FALSE	3/14/2022	9:57:50	12:26:00	148.2	55.6875	2968.6500	2	340	173	16.0	10.31	1224	510	2500	0.000	RP	3.14	2.722	0.392	5.46	2.80	0.22	3.55	1.52	3.077	17.52	3.14	0.044	3.109	0.218	5.46	0.025	5.468	0.125	5.30	3.14	0.031	3.109	0.156	5.46	0.018	5.468	0.089	5.30	117640	725	65.6	3	7.46	254	7.6	3	3.40	3899	0.095	15	0.315	0	C	B	A	20.1	967.7	48.5	445.5	swisstopo	5703006202		0317_20220314_0946_Leica GPCL3_78911 (swisstopo 39)_swisstopo.mat	none
340	Leica GPCL3	78911 (swisstopo 39)	swisstopo	FALSE	3/23/2023	16:31:17	18:52:27	141.2	55.6875	2968.6500	2	340	173	15.9	10.25	1224	510	2500	0.000	RoPr	-15.65	4.940	0.427	5.16	2.82	0.24	3.79	-2.35	4.336	24.17	-15.71	0.069	-15.708	0.210	5.16	0.039	5.164	0.120	8.03	-15.70	0.049	-15.716	0.159	5.16	0.028	5.162	0.091	8.07	117640	1065	63.6	3	8.59	224	8.2	3	4.16	4353	0.045	15	0.039	16	C	B	A	20.0	952.3	49.5	605.3	swisstopo	5703006666		0340_20230323_1620_Leica GPCL3_78911 (swisstopo 39)_swisstopo.mat	none
342	Leica GPCL3	78911 (swisstopo 39)	swisstopo	FALSE	3/24/2023	9:00:46	11:22:02	141.3	55.6875	2968.6500	2	340	173	15.9	10.28	1224	510	2500	0.000	RoPr	-6.60	4.623	0.424	5.03	2.85	0.24	3.80	-2.24	4.326	24.38	-6.67	0.055	-6.610	0.216	5.01	0.031	5.033	0.123	6.43	-6.67	0.039	-6.599	0.156	5.01	0.022	5.032	0.089	6.44	117640	1032	64.0	3	8.62	246	8.3	3	4.14	4436	0.048	15	0.043	5	C	B	A	20.0	951.7	49.5	470.0	swisstopo	5703006666		0342_20230324_0849_Leica GPCL3_78911 (swisstopo 39)_swisstopo.mat	none
362	Leica GPCL3	78911 (swisstopo 39)	swisstopo	FALSE	2/23/2024	9:46:49	13:15:13	208.4	55.6875	2968.6500	2	340	173	15.9	10.27	1224	510	2500	0.000	RoPr	-15.10	3.526	0.423	6.18	3.08	0.24	3.82	-2.77	4.325	23.96	-15.09	0.056	-15.063	0.208	6.19	0.032	6.155	0.119	6.65	-15.09	0.040	-15.064	0.153	6.19	0.023	6.156	0.087	6.66	117640	1148	64.9	3	8.83	435	7.8	3	4.10	4790	0.040	15	0.031	1	C	B	A	20.2	940.1	49.6	432.4	swisstopo	5703006666 (2024)	No inital abnormalities visible	0362_20240223_0934_Leica GPCL3_78911 (swisstopo 39)_swisstopo.mat	none
445	Leica GPCL3	78911 (swisstopo 39)	Bundesamt für Landestopografie swisstopo	FALSE	5/8/2025	14:41:59	17:37:06	175.1	55.6875	2968.6500	2	340	173	15.9	10.27	1224	540	2500	0.000	RoPr	-17.68	2.381	0.402	6.24	3.05	0.23	3.66	0.71	3.634	21.26	-17.63	0.057	-17.627	0.222	6.17	0.033	6.175	0.127	6.84	-17.62	0.040	-17.616	0.162	6.17	0.023	6.174	0.093	6.86	117640	1434	66.1	3	8.14	1094	7.8	3	4.55	3661	0.060	15	0.046	0	C	A	A	19.7	951.9	50.7	452.6	Bundesamt für Landestopografie swisstopo	9950001277	small code defects, scratches @ 73, 967 & 1370MM.	0445_20250508_1429_Leica GPCL3_78911 (swisstopo 39)_Bundesamt für Landestopografie swisstopo.mat	Small code defects, scratches @ 73, 967 & 1370mm.
					-7.63	±9.387			5.46	±0.548

294	Leica GPCL3	78987 (swisstopo 40)	swisstopo	FALSE	4/26/2021	13:26:19	15:51:04	144.8	55.6875	2968.6500	2	340	173	15.9	10.25	1224	400	2500	0.000	RP	6.58	3.576	0.324	1.99	0.22	0.19	2.92	-0.61	3.363	18.70	6.52	0.037	6.504	0.144	2.00	0.021	2.024	0.082	4.33	6.51	0.026	6.514	0.127	2.01	0.015	2.017	0.073	4.44	117640	431	62.3	3	6.75	101	8.9	3	3.48	5175	0.097	15	0.280	0	B	A	A	20.0	952.5	50.2	440.1	swisstopo	Vergleich TUM	ungebrauchte Latte, Verglcieh mit TUM	0294_20210426_1313_Leica GPCL3_78987 (swisstopo 40)_swisstopo.mat	none
326	Leica GPCL3	78987 (swisstopo 40)	swisstopo	FALSE	3/16/2022	13:56:10	16:23:33	147.4	55.6875	2968.6500	2	340	173	15.9	10.30	1224	400	2500	0.000	RP	7.79	2.308	0.317	2.72	0.67	0.18	2.87	-2.07	3.207	17.63	7.75	0.033	7.763	0.143	2.74	0.019	2.725	0.082	3.99	7.75	0.024	7.762	0.111	2.74	0.014	2.727	0.064	4.05	117640	413	66.6	3	7.31	248	7.9	3	3.39	3504	0.073	15	0.194	2	C	A	A	20.2	960.9	49.0	469.6	swisstopo	5703006202		0326_20220316_1342_Leica GPCL3_78987_swisstopo.mat	none
339	Leica GPCL3	78987 (swisstopo 40)	swisstopo	FALSE	3/23/2023	13:16:16	15:36:19	140.0	55.6875	2968.6500	2	340	173	15.9	10.25	1224	400	2500	0.000	RoPr	-96.96	3.121	0.366	2.39	1.08	0.21	3.30	-7.35	4.701	26.16	-97.00	0.043	-97.000	0.147	2.44	0.025	2.430	0.084	5.12	-97.01	0.031	-97.001	0.124	2.45	0.018	2.421	0.071	5.19	117640	834	62.8	3	8.88	463	8.4	3	4.22	4730	0.049	15	0.046	6	C	B	A	20.1	952.8	49.5	637.2	swisstopo	5703006666		0339_20230323_1304_Leica GPCL3_78987 (swisstopo 40)_swisstopo.mat	none
368	Leica GPCL3	78987 (swisstopo 40)	swisstopo	FALSE	2/28/2024	16:43:28	19:38:05	174.6	55.6875	2968.6500	2	340	173	15.9	10.26	1224	400	2500	0.000	RoPr	-86.82	2.559	0.362	3.19	0.60	0.21	3.26	-8.14	5.000	27.30	-86.90	0.034	-86.838	0.147	3.19	0.019	3.181	0.084	4.03	-86.90	0.024	-86.843	0.107	3.19	0.014	3.179	0.061	4.05	117640	1492	64.7	3	10.13	272	8.4	3	4.33	5694	0.044	15	0.037	1	C	B	A	20.1	959.7	49.4	524.3	swisstopo	5703006666 (2024)		0368_20240228_1630_Leica GPCL3_78987 (swisstopo 40)_swisstopo.mat	none
446	Leica GPCL3	78987 (swisstopo 40)	Bundesamt für Landestopografie swisstopo	FALSE	5/9/2025	10:58:46	13:53:00	174.2	55.6875	2968.6500	2	340	173	15.9	10.27	1224	400	2400	0.000	RoPr	-83.37	2.884	0.354	3.51	0.76	0.20	3.17	-4.70	3.985	22.22	-83.41	0.040	-83.413	0.153	3.52	0.023	3.523	0.088	4.67	-83.42	0.028	-83.416	0.110	3.52	0.016	3.526	0.063	4.67	117640	891	64.8	3	8.99	184	7.7	3	3.74	4655	0.042	15	0.040	10	C	B	A	19.7	955.9	50.6	457.5	Bundesamt für Landestopografie swisstopo	9950001277	Small defects/scratches @ 60, 644, 1300 & 1590mm.	0446_20250509_1047_Leica GPCL3_78987_Bundesamt für Landestopografie swisstopo.mat	Small defects/scratches @ 60, 644, 1300 & 1590mm.
					-50.61	±47.359			2.78	±0.534

164	Leica GWCL92	swisstopo 11	swisstopo	TRUE	1/21/2020	15:32:19	16:17:39	45.3	55.6875	918.3375	2	105	173	15.9	10.25	1224	207	722	4.000	RP	-386.32	5.350	0.449	3.23	12.54	0.87	2.12	-19.24	7.671	40.19	-386.29	0.099	-386.397	0.152	3.03	0.194	3.310	0.298	6.25	-386.29	0.070	-386.396	0.110	3.04	0.137	3.309	0.216	6.25	36330	438	73.7	3	8.47	102	5.6	3	2.74	1965	0.017	15	0.014	2	C	B	A	20.2	975.3	49.2	494.7	swisstopo	5703005198	Messung mit Offset, da Latte keinen magnetischen Fuss hat, Messung 0163 abgebrochen wegen zu grossem Nullpunktsoffset. Messung 0164 ist "Kontrollmessung" mit erweitertem Suchfeld (ca. 1.22 statt 0.61mm)	0164_20200121_1516_Leica GWCL92_swisstopo 11_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
166	Leica GWCL92	swisstopo 11	swisstopo	TRUE	1/22/2020	13:42:47	14:28:55	46.1	55.6875	918.3375	2	105	173	15.9	10.26	1224	175	819	4.000	RP	-389.77	4.156	0.443	3.33	11.24	0.86	2.11	-19.09	7.602	39.49	-389.84	0.113	-389.816	0.158	3.59	0.220	3.424	0.308	7.17	-389.84	0.080	-389.818	0.112	3.61	0.155	3.424	0.219	7.16	36330	415	73.8	3	8.33	108	5.7	3	2.79	1814	0.015	15	0.014	1	C	B	A	20.3	971.0	49.2	505.5	swisstopo	5703005198	Messung mit erweitertem Suchfeld wie bei 0164	0166_20200122_1327_Leica GWCL92_swisstopo 11_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
170	Leica GWCL92	swisstopo 11	swisstopo	TRUE	1/24/2020	10:02:30	10:45:54	43.4	55.6875	918.3375	2	105	173	15.9	10.27	1224	175	815	4.000	RP	-391.80	5.242	0.424	3.75	12.28	0.82	1.99	-16.59	8.351	42.90	-392.09	0.126	-391.864	0.172	3.97	0.245	3.876	0.336	7.83	-392.09	0.089	-391.864	0.138	3.98	0.173	3.876	0.270	7.84	36330	459	69.6	3	8.03	71	6.2	3	3.10	2765	0.016	15	0.011	0	C	B	A	20.4	964.5	49.1	477.7	swisstopo	5703005198	Wiederholungsmessung nach Neumontage und Sicherung Lattenfuss, Nullpunkt vermutlich stark verändert. Messung mit erweitertem Suchfeld (120px)	0170_20200124_0946_Leica GWCL92_swisstopo 11_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Fixation of custom staff foot secured with chemical threadlocker and some deformations removed
	Mittelwert	swisstopo 11					-389.41	±2.387			3.54	±0.387

167	Leica GWCL182	swisstopo 12	swisstopo	TRUE	1/22/2020	15:37:29	17:04:07	86.6	55.6875	1825.5375	2	201	173	15.9	10.29	1224	254	1549	4.000	RP	-179.85	5.548	0.483	-3.55	3.95	0.47	3.27	1.40	7.841	38.33	-179.71	0.075	-179.859	0.314	-3.39	0.073	-3.527	0.305	6.70	-179.71	0.053	-179.858	0.222	-3.39	0.052	-3.529	0.216	6.70	69546	652	71.2	3	7.24	442	6.9	3	3.09	7453	0.016	15	0.011	1	C	B	A	20.3	970.3	49.2	515.5	swisstopo	5703005198	Messung mit erweitertem Suchfeld wie bei 0164	0167_20200122_1519_Leica GWCL182_swisstopo 12_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
168	Leica GWCL182	swisstopo 12	swisstopo	TRUE	1/23/2020	9:54:18	11:19:07	84.8	55.6875	1825.5375	2	201	173	15.9	10.29	1224	254	1547	4.000	RP	-185.00	4.129	0.462	-3.27	4.00	0.45	3.12	2.77	7.712	36.09	-185.09	0.072	-185.004	0.301	-3.11	0.070	-3.248	0.292	6.39	-185.09	0.051	-185.007	0.213	-3.11	0.049	-3.245	0.207	6.40	69546	514	76.1	3	7.14	535	7.8	3	3.21	7216	0.017	15	0.012	4	C	B	A	20.4	968.4	49.1	460.8	swisstopo	5703005198		0168_20200123_0939_Leica GWCL182_swisstopo 12_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Fixation of custom staff foot secured with chemical threadlocker and some deformations removed
	Mittelwert	swisstopo 12					-182.40	±2.691			-3.25	±0.143

165	Zeiss GWLD92	swisstopo 13	swisstopo	TRUE	1/22/2020	10:38:52	11:15:21	36.5	52.7500	907.7500	2	77	173	15.9	10.25	1224	276	655	4.000	RP	187.90	5.426	0.768	3.52	3.39	1.37	2.92	-20.41	8.886	48.57	187.71	0.111	188.016	0.343	3.76	0.198	3.394	0.611	5.65	187.70	0.079	187.877	0.252	3.76	0.140	3.601	0.451	5.64	26642	795	71.3	3	9.19	161	7.4	3	4.71	639	0.017	15	0.016	2	C	B	A	20.3	972.7	49.1	498.0	swisstopo	5703005198	Messung mit erweitertem Suchfeld wie bei 0164	0165_20200122_1022_Zeiss GWLD92_swisstopo 13_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
169	Zeiss GWLD92	swisstopo 13	swisstopo	TRUE	1/23/2020	16:19:16	16:52:48	33.5	57.7500	915.0000	2	77	173	15.9	10.26	1224	230	700	4.000	RP	188.70	5.153	0.760	3.05	2.46	1.34	2.84	-18.44	8.893	47.59	188.72	0.111	188.595	0.384	3.11	0.195	3.304	0.672	5.54	188.73	0.078	188.595	0.273	3.10	0.138	3.304	0.477	5.54	26642	409	73.8	3	8.74	79	7.2	3	2.95	1485	0.023	15	0.067	0	X	B	A	20.2	964.8	49.2	490.5	swisstopo	5703005198		0169_20200123_1604_Zeiss GWLD92_swisstopo 13_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
171	Zeiss GWLD92	swisstopo 13	swisstopo	TRUE	1/24/2020	11:29:54	12:03:13	33.3	57.7500	915.0000	2	77	173	15.9	10.27	1224	230	700	4.000	RP	179.89	6.389	0.739	3.02	2.79	1.30	2.71	-16.58	9.040	47.04	179.87	0.129	179.810	0.364	3.31	0.225	3.212	0.636	6.29	179.88	0.091	179.810	0.257	3.30	0.159	3.212	0.449	6.29	26642	475	70.9	3	9.15	53	9.0	3	3.12	2325	0.017	15	0.012	0	C	B	A	20.3	964.2	49.1	490.1	swisstopo	5703005198	Wiederholungsmessung nach Neumontage und Sicherung Lattenfuss, Nullpunkt vermutlich stark verändert. Messung mit erweitertem Suchfeld (120px)	0171_20200124_1113_Zeiss GWLD92_swisstopo 13_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. Fixation of custom staff foot secured with chemical threadlocker and some deformations removed
173	Zeiss GWLD92	swisstopo 13	swisstopo	TRUE	1/24/2020	14:11:52	14:45:00	33.1	57.7500	915.0000	2	77	173	15.9	10.27	1224	230	700	4.000	RP	188.56	5.277	0.752	2.53	2.70	1.32	2.75	-16.55	9.065	47.45	188.52	0.124	188.548	0.376	2.67	0.217	2.642	0.659	6.05	188.51	0.087	188.548	0.265	2.69	0.153	2.642	0.463	6.04	26642	501	70.9	3	9.22	43	9.0	3	3.08	2488	0.017	15	0.012	0	C	B	A	20.3	962.2	49.2	491.9	swisstopo	5703005198	Wiederholungsmessung nach Neumontage und Sicherung Lattenfuss, Nullpunkt vermutlich stark verändert. Messung mit erweitertem Suchfeld (120px) Wiederholung von Messung 0171 wegen "Schere" in Hin- und Rückmessung. 0172 manuell abgebrochen wgen Positionierungsfehler, Latte nicht neu ausgerichtet	0173_20200124_1356_Zeiss GWLD92_swisstopo 13_swisstopo.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing.
	Mittelwert	swisstopo 13					186.21	±3.673			3.21	±0.386

198	Leica GPCL3	30155	Amt für Raumentwicklung, Kanton Zürich	TRUE	3/10/2020	9:25:59	11:51:48	145.8	55.6875	2968.6500	2	340	173	15.9	10.27	1224	270	2720	0.000	RP	16.47	4.963	0.400	3.34	1.29	0.23	3.45	11.27	7.334	38.82	16.33	0.055	16.428	0.229	3.32	0.031	3.360	0.131	6.27	16.34	0.039	16.430	0.162	3.31	0.022	3.359	0.093	6.27	117640	2289	67.8	3	8.37	138	5.7	3	2.59	8885	0.108	15	0.175	5	C	B	A	20.2	957.7	49.4	437.7	Amt für Raumentwicklung, Kanton Zürich	Fixpunkte/83210.09.003	Lattenlagerung sehr weit aussen, aber symmetrisch, Latte ev. leicht durchgebogen. Visuell keine groben Schäden an Gehäuse oder Code feststellbar. Lattenstativ lässt sich nicht entfernen. "Kopfschraube"rechts dreht durch (vermutlich Gewindehülse beschädigt und dreht durch)	0198_20200310_0908_Leica GPCL3_30155_Amt für Raumentwicklung, Kanton Zürich.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. The index correction refers to the staff base plate and it is determined by a reference with three point support magnetically attached to the base. The tested staff shows local signs of wear with a few 0.01mm deviation. Such local deviations cannot be determined with the calibration facility! Levelling staff bipod cannot be removed. Mounting screw on the right side turns through without loosening.
356	Leica GPCL3	30155	Amt für Raumentwicklung Zürich	FALSE	5/15/2023	10:19:42	12:55:41	156.0	55.6875	2968.6500	2	340	173	15.9	10.25	1224	500	2500	0.000	RoPr	9.03	5.189	0.374	3.46	1.34	0.21	3.34	12.16	6.660	34.65	8.87	0.055	8.883	0.260	3.43	0.032	3.498	0.149	6.54	8.86	0.039	8.883	0.186	3.43	0.022	3.498	0.107	6.56	117640	1253	67.9	3	6.41	941	5.3	3	3.45	3137	0.324	15	0.763	0	C	B	A	19.7	954.1	50.2	448.5	Amt für Raumentwicklung Zürich	-		0356_20230515_1008_Leica GPCL3_30155_Amt für Raumentwicklung Zürich.mat	Levelling staff bipod cannot be removed. Mounting screw on the right side turns through without loosening. Feet had to be removed individually.
					12.60	±3.737			3.37	±0.058

200	Leica GPCL3	30152	Amt für Raumentwicklung, Kanton Zürich	TRUE	3/10/2020	13:36:32	16:00:59	144.4	55.6875	2968.6500	2	340	173	15.9	10.29	1224	500	2500	0.000	RP	3.15	4.134	0.433	5.23	3.05	0.25	3.89	12.65	6.695	36.39	3.22	0.054	3.108	0.263	5.12	0.031	5.252	0.151	6.37	3.22	0.038	3.109	0.187	5.12	0.022	5.252	0.107	6.38	117640	1635	64.8	3	7.28	352	7.0	3	4.14	3703	0.104	15	0.131	1	C	B	A	20.2	957.0	49.3	446.3	Amt für Raumentwicklung, Kanton Zürich	Fixpunkte/83210.09.003	lokale Abnutzung am Lattenfusspunkt. Schaden im Lattengehäuse bei 1.90m (links aussen). kleiner Schaden im Lattengehäuse bei 1.325m (rechts aussen). ausgeschlagene Kante bei 498.xxxmm. grober Schlag im Codeband (nicht auf Kante bei 1707mm)	0200_20200310_1321_Leica GPCL3_30152_Amt für Raumentwicklung, Kanton Zürich.mat	Data reprocessed on 06.07.2020 with updated software; difference of estimated parameters with respect to original processing is not significant, STD of estimated parameters is lower with new processing. The index correction refers to the staff base plate and it is determined by a reference with three point support magnetically attached to the base. The tested staff shows local signs of wear with a few 0.01mm deviation. Such local deviations cannot be determined with the calibration facility! At approximately 1.70m the code-tape shows a dent with a depth of up to 0.9mm. Although none of the code-edges is directly damaged, the local deformation of the code-tape is clearly visible in the calibration results.
354	Leica GPCL3	30152	Amt für Raumentwicklung Zürich	FALSE	5/12/2023	8:07:51	10:38:17	150.4	55.6875	2968.6500	2	340	173	15.9	10.27	1224	500	2500	0.000	RoPr	10.10	4.188	0.443	4.88	3.26	0.25	3.94	12.77	6.265	33.60	10.00	0.058	9.805	0.297	4.79	0.033	5.027	0.170	6.74	9.99	0.041	9.799	0.211	4.80	0.023	5.024	0.121	6.74	117640	1601	66.4	3	6.37	910	5.9	3	3.40	3491	0.216	15	0.551	0	C	B	A	19.7	953.6	49.9	462.1	Amt für Raumentwicklung Zürich	-	Lattenstativ konnte nicht vollständig entfernt werden (Schrauben der Kopfplatte "festgefressen"). Beide Füsse separat demontiert, da sonst eine Kalibrierung nicht möglich wäre.	0354_20230512_0756_Leica GPCL3_30152_Amt für Raumentwicklung Zürich.mat	Levelling staff bipod can not be removed. Screws are stuck. Feet had to be removed individually. At approximately 1.70m the code shows a dent with a depth of about 0.9mm. The local deformation is clearly visible in the calibration results.
					6.61	±3.385			4.96	±0.162

250	Leica GPCL3	72774	Stadt Zürich, Geomatik + Vermessung	FALSE	9/21/2020	10:57:53	13:21:14	143.3	55.6875	2968.6500	2	340	173	16.0	10.31	1224	600	2400	0.000	RP	-43.25	4.897	0.363	3.01	0.92	0.21	3.23	-26.68	8.849	43.07	-43.21	0.048	-43.265	0.143	2.99	0.028	3.008	0.082	5.70	-43.22	0.035	-43.272	0.131	3.00	0.020	3.012	0.075	5.80	117640	1132	65.3	3	6.26	759	10.3	3	4.72	3641	0.054	15	0.057	12	X	C	A	19.7	956.6	50.1	419.3	Stadt Zürich, Geomatik + Vermessung	-		0250_20200921_1046_Lieca GPCL3_72774_Stadt Zürich, Geomatik + Vermessung.mat	none
260	Leica GPCL3	72774	Stadt Zürich, Geomatik + Vermessung	FALSE	10/7/2020	13:46:53	16:10:20	143.5	55.6875	2968.6500	2	340	173	15.9	10.30	1224	600	2400	0.000	RP	-47.95	4.848	0.369	3.85	1.02	0.21	3.28	-26.62	8.817	42.79	-47.96	0.048	-47.985	0.140	3.85	0.027	3.866	0.080	5.62	-47.96	0.034	-47.965	0.107	3.85	0.019	3.855	0.061	5.65	117640	1133	66.7	3	6.32	857	10.3	3	3.98	6059	0.051	15	0.038	1	X	C	A	19.7	958.5	49.7	423.4	Stadt Zürich, Geomatik + Vermessung	-	Cal. No. 259 manually aborted due to too high temperature effect between forward- and backward run. Calibration 260 started without change of alignment	0260_20201007_1331_Leica GPCL3_72774_Stadt Zürich, Geomatik + Vermessung.mat	none
397	Leica GPCL3	72774	Stadt Zürich, Geomatik + Vermessung	FALSE	3/31/2025	15:09:41	18:05:19	175.6	55.6875	2968.6500	2	340	173	15.9	10.26	1224	600	2400	0.000	RoPr	-49.18	4.592	0.342	3.48	1.23	0.20	3.07	-26.35	8.451	39.48	-49.14	0.047	-49.205	0.133	3.48	0.027	3.486	0.076	5.54	-49.14	0.033	-49.205	0.095	3.48	0.019	3.486	0.054	5.54	117640	1164	64.1	3	5.50	1126	8.3	3	4.22	3299	0.062	15	0.055	0	X	C	A	19.8	962.0	50.6	451.5	Stadt Zürich, Geomatik + Vermessung	-	General signs of wear, but no strong defects. / Small damages to the code e.g. @ 103 to 110 cm, 170 & 290 cm.	0397_20250331_1458_Leica GPCL3_72774_Stadt Zürich, Geomatik + Vermessung.mat	General signs of wear, but no strong defects. Small damages to the code e.g. @ 103 to 110 cm, 170 & 290 cm.
					-46.78	±2.559			3.44	±0.350

251	Leica GPCL3	72770	Stadt Zürich, Geomatik + Vermessung	FALSE	9/21/2020	14:35:23	17:00:10	144.8	55.6875	2968.6500	2	340	173	15.9	10.25	1224	500	2330	0.000	RP	-72.68	4.762	0.419	3.44	3.32	0.24	3.76	-28.66	9.099	43.35	-72.83	0.069	-72.800	0.165	3.59	0.039	3.529	0.094	8.14	-72.81	0.049	-72.791	0.145	3.58	0.028	3.524	0.083	8.22	117640	1169	64.0	3	6.46	847	9.7	3	4.79	3176	0.046	15	0.028	1	X	C	A	19.7	955.2	50.0	411.2	Stadt Zürich, Geomatik + Vermessung	-		0251_20200921_1416_Lieca GPCL3_72770_Stadt Zürich, Geomatik + Vermessung.mat	none
261	Leica GPCL3	72770	Stadt Zürich, Geomatik + Vermessung	FALSE	10/7/2020	17:22:05	19:44:55	142.8	55.6875	2968.6500	2	340	173	15.9	10.29	1224	500	2330	0.000	RP	-71.17	4.598	0.424	3.23	3.92	0.24	3.77	-28.88	9.168	43.78	-71.39	0.068	-71.348	0.174	3.43	0.039	3.338	0.100	7.92	-71.35	0.049	-71.215	0.191	3.39	0.028	3.274	0.109	8.19	117640	1298	64.8	3	6.51	954	9.6	3	3.98	5968	0.056	15	0.049	1	X	C	A	19.8	960.1	49.3	415.0	Stadt Zürich, Geomatik + Vermessung	-		0261_20201007_1709_Leica GPCL3_72770_Stadt Zürich, Geomatik + Vermessung.mat	none
262	Leica GPCL3	72770	Stadt Zürich, Geomatik + Vermessung	FALSE	10/8/2020	8:50:58	11:12:40	141.7	55.6875	2968.6500	2	340	173	15.9	10.29	1224	500	2330	0.000	RP	-68.95	4.548	0.429	2.75	3.90	0.25	3.81	-28.88	9.155	43.71	-69.13	0.068	-69.109	0.183	2.91	0.039	2.852	0.105	7.94	-69.14	0.048	-69.111	0.134	2.92	0.028	2.854	0.077	7.96	117640	1331	64.0	3	6.42	1008	9.4	3	3.97	5711	0.063	15	0.054	3	X	C	A	19.8	963.7	48.9	421.5	Stadt Zürich, Geomatik + Vermessung	-	0261 showed high differences between forward and backward run. 0262 startet with only minor adjustments to the staff.	0262_20201008_0839_Leica GPCL3_72770_Stadt Zürich, Geomatik + Vermessung.mat	none
395	Leica GPCL3	72770	Stadt Zürich, Geomatik + Vermessung	FALSE	3/28/2025	14:27:12	17:22:48	175.6	55.6875	2968.6500	2	340	173	15.9	10.27	1224	475	2330	0.000	RoPr	-75.04	4.220	0.389	3.83	3.83	0.22	3.51	-27.82	8.649	39.56	-75.08	0.074	-75.135	0.160	3.88	0.043	3.893	0.092	8.83	-75.09	0.053	-75.135	0.124	3.88	0.030	3.894	0.071	8.85	117640	1252	63.0	3	5.01	1043	7.2	3	3.97	2534	0.054	15	0.044	2	X	C	A	19.9	949.8	50.6	469.0	Stadt Zürich, Geomatik + Vermessung	-	Scratches @ 185.5 & 264 to 268cm. / Sticky dirt at ~20cm removed. / Both top screws for staff-tripod were loose. / Minor dents in right housing side @ 87 and 248cm.	0395_20250328_1414_Leica GPCL3_72770_Stadt Zürich, Geomatik + Vermessung.mat	Scratches @ 185.5 & 264 to 268cm. Sticky dirt at ~20cm removed. Both top screws for staff-tripod were loose. Minor dents in right housing side @ 87 and 248cm.
396	Leica GPCL3	72770	Stadt Zürich, Geomatik + Vermessung	FALSE	3/31/2025	9:34:20	12:30:04	175.7	55.6875	2968.6500	2	340	173	15.9	10.26	1224	475	2330	0.000	RoPr	-72.45	4.199	0.391	3.57	3.83	0.22	3.51	-27.77	8.620	39.43	-72.51	0.074	-72.561	0.168	3.64	0.042	3.645	0.096	8.79	-72.52	0.052	-72.563	0.125	3.64	0.030	3.646	0.071	8.81	117640	1280	62.9	3	5.00	1062	7.3	3	3.98	2540	0.059	15	0.047	1	X	C	A	19.8	961.9	50.5	444.4	Stadt Zürich, Geomatik + Vermessung	-	Scratches @ 185.5 & 264 to 268cm. / Sticky dirt at ~20cm removed. / Both top screws for staff-tripod loose. / Minor dents in right housing side @ 87 and 248cm.	0396_20250331_0922_Leica GPCL3_72770_Stadt Zürich, Geomatik + Vermessung.mat	Scratches @ 185.5 & 264 to 268cm. Sticky dirt at ~20cm removed. Both top screws for staff-tripod loose. Minor dents in right housing side @ 87 and 248cm.
398	Leica GPCL3	72770	Stadt Zürich, Geomatik + Vermessung	FALSE	4/1/2025	11:06:11	14:01:53	175.7	55.6875	2968.6500	2	340	173	15.9	10.26	1224	475	2330	0.000	RoPr	-81.07	4.190	0.393	3.50	3.81	0.23	3.54	-27.80	8.650	39.71	-81.12	0.074	-81.161	0.162	3.56	0.042	3.566	0.093	8.80	-81.12	0.052	-81.161	0.116	3.56	0.030	3.566	0.067	8.80	117640	1183	62.9	3	5.03	1070	7.2	3	3.99	2509	0.050	15	0.038	0	X	C	A	19.8	957.6	50.6	453.0	Stadt Zürich, Geomatik + Vermessung	-	Scratches @ 185.5 & 264 to 268cm. / Sticky dirt at ~20cm removed. / Both top screws for staff-tripod loose. / Minor dents in right housing side @ 87 and 248cm.	0398_20250401_1055_Leica GPCL3_72770_Stadt Zürich, Geomatik + Vermessung.mat	Scratches @ 185.5 & 264 to 268cm. Sticky dirt at ~20cm removed. Both top screws for staff-tripod loose. Minor dents in right housing side @ 87 and 248cm.
					-73.67	±3.774			3.49	±0.295

256	Leica GPCL2	76191	Stadt Zürich, Geomatik + Vermessung	FALSE	9/22/2020	16:11:27	17:44:40	93.2	55.6875	1934.8875	2	216	173	15.9	10.29	1224	525	1470	0.000	RP	-21.20	9.430	0.496	2.71	7.84	0.44	3.58	-42.85	9.309	45.49	-21.86	0.083	-21.120	0.267	2.76	0.075	2.680	0.239	7.88	-21.86	0.059	-21.139	0.194	2.76	0.053	2.684	0.174	7.90	74736	788	64.7	3	6.98	381	11.3	3	4.41	5820	0.031	15	0.016	1	X	C	A	19.6	951.3	50.1	416.8	Stadt Zürich, Geomatik + Vermessung	-		0256_20200922_1600_Leica GPCL2_76191_Stadt Zürich, Geomatik + Vermessung.mat	none
263	Leica GPCL2	76191	Stadt Zürich, Geomatik + Vermessung	FALSE	10/8/2020	13:22:44	14:53:28	90.7	55.6875	1934.8875	2	216	173	15.9	10.24	1224	525	1470	0.000	RP	-25.86	8.249	0.497	3.02	7.70	0.45	3.57	-42.94	9.268	45.20	-26.14	0.083	-25.830	0.262	3.01	0.074	3.036	0.235	7.80	-26.15	0.058	-25.831	0.185	3.01	0.052	3.036	0.166	7.80	74736	801	63.0	3	6.63	436	10.9	3	4.40	5626	0.039	15	0.033	3	X	C	A	19.7	962.2	49.3	410.7	Stadt Zürich, Geomatik + Vermessung	-		0263_20201008_1310_Leica GPCL2_76191_Stadt Zürich, Geomatik + Vermessung.mat	none
392	Leica GPCL2	76191	Stadt Zürich, Geomatik + Vermessung	FALSE	3/25/2025	10:46:09	12:37:56	111.8	55.6875	1934.8875	2	216	173	15.9	10.26	1224	600	1425	0.000	RoPr	-74.33	5.627	0.427	3.38	3.91	0.38	3.06	-43.07	8.889	41.70	-74.44	0.068	-74.359	0.277	3.42	0.061	3.398	0.248	6.41	-74.44	0.048	-74.379	0.198	3.42	0.043	3.403	0.178	6.43	74736	655	64.5	3	5.40	1188	8.1	3	3.79	3993	0.045	15	0.041	0	X	C	A	19.8	958.8	50.7	471.8	Stadt Zürich, Geomatik + Vermessung	-	Heavy signs of wear. / Small defect in housing @175cm (front/right).	0392_20250325_1034_Leica GPCL2_76191_Stadt Zürich, Geomatik + Vermessung.mat	Heavy signs of wear. Small defect in housing @175cm (front/right).
393	Leica GPCL2	76191	Stadt Zürich, Geomatik + Vermessung	FALSE	3/25/2025	15:22:19	17:13:50	111.5	55.6875	1934.8875	2	216	173	15.9	10.26	1224	600	1425	0.000	RoPr	-66.52	5.348	0.424	3.25	3.93	0.38	3.04	-43.15	8.905	41.52	-66.55	0.067	-66.594	0.268	3.31	0.060	3.305	0.240	6.38	-66.54	0.048	-66.595	0.189	3.30	0.043	3.305	0.170	6.38	74736	670	64.6	3	5.39	1262	8.2	3	3.79	3971	0.046	15	0.037	1	X	C	A	19.8	958.1	50.7	463.6	Stadt Zürich, Geomatik + Vermessung	-	Heavy signs of wear. / Small defect in housing @175cm (front/right).	0393_20250325_1511_Leica GPCL2_76191_Stadt Zürich, Geomatik + Vermessung.mat	Heavy signs of wear. Small defect in housing @175cm (front/right).
					-47.25	±23.460			3.12	±0.258

257	GeoZ Midi	Nr. 10	Stadt Zürich, Geomatik + Vermessung	FALSE	10/2/2020	10:03:31	10:49:15	45.7	45.5625	934.5375	2	107	171	15.9	10.40	1224	155	680	0.000	RP	2.72	2.828	0.466	-5.57	2.65	0.91	2.26	17.77	4.401	24.22	2.75	0.064	2.510	0.240	-5.71	0.125	-4.869	0.482	4.12	2.76	0.045	2.518	0.173	-5.73	0.089	-4.885	0.347	4.12	36594	624	77.9	3	5.78	412	6.2	3	3.42	644	0.016	15	0.010	6	C	B	A	19.8	933.4	49.9	423.0	Stadt Zürich, Geomatik + Vermessung	-	very difficult to align coaxiality	0257_20201002_0951_GeoZ Midi_Nr. 10_Stadt Zürich, Geomatik + Vermessung.mat	none
427	GeoZ Midi (0.9m)	GeoZ 6-10	Stadt Zürich, Geomatik + Vermessung	FALSE	4/14/2025	17:19:10	18:13:42	54.5	45.5625	938.5875	2	108	173	15.9	10.26	1224	155	680	0.000	RoPr	7.47	4.117	0.505	-5.48	4.44	0.99	2.52	15.50	4.792	26.27	7.48	0.085	7.280	0.243	-5.49	0.167	-4.757	0.487	5.64	7.49	0.060	7.280	0.173	-5.49	0.118	-4.757	0.345	5.64	37368	857	73.0	3	6.30	576	5.6	3	3.74	803	0.040	15	0.056	0	C	B	A	19.6	946.0	51.0	466.2	Stadt Zürich, Geomatik + Vermessung	-	none	0427_20250414_1707_GeoZ Midi (0.9m)_GeoZ 6-10_Stadt Zürich, Geomatik + Vermessung.mat	Possibly a bit of dirt in the gcode guide (audible noise).
429	GeoZ Midi (0.9m)	GeoZ 6-10	Stadt Zürich, Geomatik + Vermessung	FALSE	4/15/2025	11:41:13	12:35:19	54.1	45.5625	934.5375	2	107	173	15.9	10.27	1224	155	680	0.000	RoPr	0.66	3.909	0.506	-5.23	2.98	0.99	2.50	15.43	4.862	26.99	0.71	0.079	0.445	0.244	-5.22	0.154	-4.508	0.489	5.17	0.70	0.056	0.446	0.179	-5.21	0.109	-4.509	0.359	5.18	37022	512	74.9	3	5.72	343	5.6	3	3.05	742	0.032	15	0.054	0	C	B	A	19.6	942.6	51.2	457.7	Stadt Zürich, Geomatik + Vermessung	-	Scratchy noises inside the code guide fins (maybe dirt). / Run with rubberized support also in the back.	0429_20250415_1130_GeoZ Midi (0.9m)_GeoZ 6-10_Stadt Zürich, Geomatik + Vermessung.mat	Scratchy noises inside the code guide fins (maybe dirt).
430	GeoZ Midi (0.9m)	GeoZ 6-10	Stadt Zürich, Geomatik + Vermessung	FALSE	4/15/2025	13:36:17	14:30:40	54.4	45.5625	934.5375	2	107	173	15.9	10.26	1224	155	680	0.000	RoPr	0.72	3.917	0.507	-4.56	2.99	0.99	2.51	15.44	4.852	26.59	0.74	0.079	0.495	0.248	-4.49	0.154	-3.812	0.497	5.19	0.74	0.056	0.551	0.172	-4.48	0.109	-3.898	0.344	5.19	37022	476	74.8	3	5.71	367	5.6	3	3.04	702	0.030	15	0.053	0	C	B	A	19.6	942.7	51.1	466.1	Stadt Zürich, Geomatik + Vermessung	-	Scratchy noises inside the code guide fins (maybe dirt). / Run with rubberized support also in the back.	0430_20250415_1325_GeoZ Midi (0.9m)_GeoZ 6-10_Stadt Zürich, Geomatik + Vermessung.mat	Scratchy noises inside the code guide fins (maybe dirt).
431	GeoZ Midi (0.9m)	GeoZ 6-10	Stadt Zürich, Geomatik + Vermessung	FALSE	4/15/2025	15:09:53	16:04:08	54.3	45.5625	934.5375	2	107	173	15.9	10.27	1224	155	680	0.000	RoPr	3.57	3.882	0.503	-4.38	3.08	0.99	2.49	15.40	4.833	26.52	3.59	0.079	3.411	0.239	-4.31	0.154	-3.755	0.479	5.17	3.59	0.056	3.363	0.181	-4.31	0.109	-3.682	0.363	5.19	37022	521	73.1	3	5.58	363	5.5	3	2.97	709	0.029	15	0.041	0	C	B	A	19.6	942.7	51.2	467.6	Stadt Zürich, Geomatik + Vermessung	-	Scratchy noises inside the code guide fins (maybe dirt).	0431_20250415_1458_GeoZ Midi (0.9m)_GeoZ 6-10_Stadt Zürich, Geomatik + Vermessung.mat	Scratchy noises inside the code guide fins (maybe dirt).
					3.05	±2.486			-5.04	±0.559

258	GeoZ Mini	Nr. 06	Stadt Zürich, Geomatik + Vermessung	FALSE	10/2/2020	14:52:52	15:03:22	10.5	45.5625	215.6625	2	26	173	16.0	10.31	1224	20	225	0.000	RP	7.37	4.455	1.764	9.48	17.65	12.73	3.03	12.62	7.084	41.37	7.39	0.230	7.410	0.359	8.62	1.662	8.961	2.585	5.43	7.40	0.163	7.410	0.274	8.58	1.178	8.961	1.974	5.44	8996	76	60.4	3	9.93	0	7.7	3	4.04	489	0.020	15	0.011	0	X	X	X	19.8	928.8	49.7	448.1	Stadt Zürich, Geomatik + Vermessung	-		0258_20201002_1440_GeoZ Mini_Nr. 06_Stadt Zürich, Geomatik + Vermessung.mat	none
434	GeoZ Mini (0.2m)	GeoZ 6-9	Stadt Zürich, Geomatik + Vermessung	FALSE	4/16/2025	14:04:59	14:17:42	12.7	45.5625	215.6625	2	26	173	15.9	10.25	1224	20	220	0.000	RoPr	-9.59	4.127	1.754	12.85	15.43	12.68	2.98	11.10	7.077	41.22	-9.87	0.200	-9.651	0.297	11.84	1.445	12.784	2.138	4.67	-9.87	0.141	-9.651	0.208	11.84	1.022	12.784	1.496	4.67	8996	59	60.4	3	9.08	0	7.0	3	3.92	572	0.041	15	0.025	0	X	X	X	19.7	948.4	51.0	459.8	Stadt Zürich, Geomatik + Vermessung	-	None.	0434_20250416_1351_GeoZ Mini (0.2m)_GeoZ 6-9_Stadt Zürich, Geomatik + Vermessung.mat	None.
435	GeoZ Mini (0.2m)	GeoZ 6-9	Stadt Zürich, Geomatik + Vermessung	FALSE	4/16/2025	14:57:11	15:09:57	12.8	45.5625	215.6625	2	26	173	15.9	10.25	1224	20	220	0.000	RoPr	-12.62	4.488	1.770	15.18	15.12	12.79	3.04	11.02	7.087	41.36	-12.83	0.204	-12.617	0.331	14.24	1.472	14.638	2.383	4.79	-12.83	0.144	-12.617	0.231	14.23	1.040	14.638	1.658	4.79	8996	45	60.5	3	9.19	1	7.0	3	3.93	547	0.045	15	0.027	0	X	X	X	19.6	948.6	51.0	463.6	Stadt Zürich, Geomatik + Vermessung	-	None.	0435_20250416_1446_GeoZ Mini (0.2m)_GeoZ 6-9_Stadt Zürich, Geomatik + Vermessung.mat	None.
439	GeoZ Mini (0.2m)	GeoZ 6-9	Stadt Zürich, Geomatik + Vermessung	FALSE	4/25/2025	10:08:13	10:20:56	12.7	45.5625	215.6625	2	26	173	15.9	10.28	1224	20	225	0.000	RoPr	-14.44	4.053	1.695	5.96	15.00	12.24	2.92	11.25	6.932	39.82	-14.63	0.205	-14.492	0.319	6.21	1.476	5.920	2.298	4.84	-14.62	0.145	-14.492	0.224	6.16	1.044	5.920	1.608	4.84	8996	44	61.5	3	9.75	0	7.5	3	4.05	487	0.050	15	0.035	0	X	X	X	19.7	957.2	50.8	438.9	Stadt Zürich, Geomatik + Vermessung	-	none	0439_20250425_0957_GeoZ Mini (0.2m)_GeoZ 6-9_Stadt Zürich, Geomatik + Vermessung.mat	None.
					-5.10	±8.919			11.55	±2.317

265	Leica GPCL3	24505 (GeoZ 7-1)	Stadt Zürich, Geomatik + Vermessung	FALSE	11/12/2020	13:11:29	15:34:25	142.9	55.6875	2968.6500	2	340	173	15.9	10.29	1224	545	2310	0.000	RP	-75.25	2.597	0.326	-4.96	0.98	0.19	2.89	10.47	4.756	27.74	-75.06	0.031	-75.302	0.211	-5.04	0.018	-4.934	0.121	3.63	-75.08	0.022	-75.302	0.161	-5.04	0.013	-4.934	0.092	3.72	117640	1639	58.1	3	4.79	2540	8.0	3	4.02	2716	0.064	15	0.041	0	C	B	A	19.9	961.8	49.4	437.4	Stadt Zürich, Geomatik + Vermessung	-	Starker Schlag auf Bandführung (rechts) bei ca. 558mm, Band aber frei beweglich. Kleine Schlag auf Bandführung (rechts) beibei ca. 2860mm, Band aber frei beweglich. "Abrieb" in der Mitte des Codes zwischen ca. 2800 und 2880mm, sowie 2540 bis 2680mm. Allgemein starke Gebrauchsspuren.	0265_20201112_1258_Leica GPCL3_24505 (GeoZ 7-1)_Stadt Zürich, Geomatik + Vermessung.mat	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free Abrasions near the code-center from appr. 2800 to 2880mm and 2540 to 2680mm Signs of wear visible throughout the whole length
266	Leica GPCL3	24505 (GeoZ 7-1)	Stadt Zürich, Geomatik + Vermessung	FALSE	11/16/2020	10:47:24	13:08:28	141.1	55.6875	2968.6500	2	340	173	15.9	10.29	1224	545	2310	0.000	RP	-75.63	6.438	0.328	-4.69	1.00	0.19	2.91	10.46	4.739	27.50	-76.20	0.032	-75.822	0.214	-4.76	0.018	-4.611	0.123	3.67	-76.19	0.023	-75.822	0.154	-4.76	0.013	-4.611	0.088	3.69	117640	1663	58.9	3	4.87	2522	8.2	3	4.05	2697	0.058	15	0.042	2	C	B	A	19.9	960.8	49.3	422.1	Stadt Zürich, Geomatik + Vermessung	-	Nachmessung zu 0265. Latte seit 12.11.2020 unverändert auf Komparator. Starker Schlag auf Bandführung (rechts) bei ca. 558mm, Band aber frei beweglich. Kleine Schlag auf Bandführung (rechts) beibei ca. 2860mm, Band aber frei beweglich. "Abrieb" in der Mitte des Codes zwischen ca. 2800 und 2880mm, sowie 2540 bis 2680mm. Allgemein starke Gebrauchsspuren.	0266_20201116_1032_Leica GPCL3_24505 (GeoZ 7-1)_Stadt Zürich, Geomatik + Vermessung.mat	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free. Abrasions near the code-center from appr. 2800 to 2880mm and 2540 to 2680mm. Signs of wear visible throughout the whole length.
267	Leica GPCL3	24505 (GeoZ 7-1)	Stadt Zürich, Geomatik + Vermessung	FALSE	11/16/2020	14:25:26	16:47:09	141.7	55.6875	2968.6500	2	340	173	15.9	10.30	1224	545	2310	0.000	RP	-76.14	2.684	0.328	-4.89	1.01	0.19	2.91	10.47	4.755	27.73	-76.02	0.032	-76.178	0.214	-4.93	0.018	-4.869	0.123	3.78	-76.01	0.023	-76.178	0.152	-4.93	0.013	-4.869	0.087	3.79	117640	1555	59.5	3	4.91	2647	8.3	3	4.08	2838	0.065	15	0.048	4	C	B	A	19.9	961.9	49.5	411.7	Stadt Zürich, Geomatik + Vermessung	-	Nachmessung zu 0266, da dort vermutlich ein Problem bei der Kamerakalibrierung aufgetreten ist. Latte seit 12.11.2020 unverändert auf Komparator. Starker Schlag auf Bandführung (rechts) bei ca. 558mm, Band aber frei beweglich. Kleine Schlag auf Bandführung (rechts) beibei ca. 2860mm, Band aber frei beweglich. "Abrieb" in der Mitte des Codes zwischen ca. 2800 und 2880mm, sowie 2540 bis 2680mm. Allgemein starke Gebrauchsspuren.	0267_20201116_1412_Leica GPCL3_24505 (GeoZ 7-1)_Stadt Zürich, Geomatik + Vermessung.mat	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free. Abrasions near the code-center from appr. 2800 to 2880mm and 2540 to 2680mm. Signs of wear visible throughout the whole length.
404	Leica GPCL3	24505 (GeoZ 7-1)	Stadt Zürich, Geomatik + Vermessung	FALSE	4/4/2025	15:12:51	18:07:34	174.7	55.6875	2967.6375	2	340	173	15.9	10.25	1224	570	2350	0.000	RoPr	-75.87	2.686	0.322	-5.13	0.94	0.18	2.84	10.36	4.694	27.50	-75.64	0.032	-75.965	0.206	-5.15	0.018	-5.067	0.118	3.75	-75.64	0.023	-75.960	0.158	-5.15	0.013	-5.070	0.090	3.84	117640	2215	60.5	3	5.19	2709	8.5	3	4.43	2770	0.061	15	0.046	1	C	B	A	19.8	955.0	50.7	492.1	Stadt Zürich, Geomatik + Vermessung	-	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free. / Abrasions near the code-center from appr. 2800 to 2910mm and 2540 to 2680mm. / Strong signs of wear.	0404_20250404_1500_Leica GPCL3_24505 (GeoZ 7-1)_Stadt Zürich, Geomatik + Vermessung.mat	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free. Abrasions near the code-center from appr. 2800 to 2910mm and 2540 to 2680mm. Strong signs of wear.
405	Leica GPCL3	24505 (GeoZ 7-1)	Stadt Zürich, Geomatik + Vermessung	FALSE	4/7/2025	8:36:37	11:31:14	174.6	55.6875	2967.6375	2	340	173	15.9	10.25	1224	570	2350	0.000	RoPr	-93.71	2.162	0.335	-4.36	0.92	0.19	2.95	10.18	4.765	28.63	-93.51	0.033	-93.805	0.220	-4.41	0.019	-4.315	0.126	3.82	-93.50	0.024	-93.732	0.170	-4.42	0.014	-4.348	0.098	3.93	117640	2048	60.5	3	5.21	2719	8.5	3	4.46	2784	0.064	15	0.054	1	C	B	A	19.8	960.4	50.7	452.4	Stadt Zürich, Geomatik + Vermessung	-	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free. / Abrasions near the code-center from appr. 2800 to 2910mm and 2540 to 2680mm. / Strong signs of wear.	0405_20250407_0823_Leica GPCL3_24505 (GeoZ 7-1)_Stadt Zürich, Geomatik + Vermessung.mat	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free. Abrasions near the code-center from appr. 2800 to 2910mm and 2540 to 2680mm. Strong signs of wear.
406	Leica GPCL3	24505 (GeoZ 7-1)	Stadt Zürich, Geomatik + Vermessung	FALSE	4/7/2025	13:46:17	16:40:27	174.2	55.6875	2967.6375	2	340	173	15.9	10.25	1224	570	2350	0.000	RoPr	-85.52	2.841	0.332	-4.40	0.97	0.19	2.93	10.27	4.750	28.22	-85.22	0.032	-85.602	0.218	-4.44	0.019	-4.361	0.125	3.78	-85.21	0.023	-85.602	0.154	-4.44	0.013	-4.361	0.089	3.78	117640	2052	60.5	3	5.22	2695	8.5	3	4.43	2724	0.061	15	0.053	1	C	B	A	19.8	960.3	50.7	502.2	Stadt Zürich, Geomatik + Vermessung	-	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free. / Abrasions near the code-center from appr. 2800 to 2910mm and 2540 to 2680mm. / Strong signs of wear.	0406_20250407_1334_Leica GPCL3_24505 (GeoZ 7-1)_Stadt Zürich, Geomatik + Vermessung.mat	Damages of the code-guide at appr. 558mm and appr. 2860mm but code-tape is still free. Abrasions near the code-center from appr. 2800 to 2910mm and 2540 to 2680mm. Strong signs of wear.
					-80.27	±6.862			-4.79	±0.281

268	Leica GPCL3	24514 (GeoZ 7-2)	Stadt Zürich, Geomatik + Vermessung	FALSE	11/17/2020	9:39:23	12:01:15	141.9	55.6875	2968.6500	2	340	173	15.9	10.29	1224	1080	2100	0.000	RP	-115.33	7.242	0.389	-5.87	3.24	0.22	3.46	4.75	6.987	39.68	-114.84	0.053	-115.338	0.210	-6.11	0.030	-5.820	0.120	6.08	-114.85	0.037	-115.344	0.156	-6.10	0.021	-5.817	0.089	6.12	117640	1646	55.8	3	5.16	1192	8.1	3	4.58	2653	0.061	15	0.040	13	C	A	A	19.8	969.3	49.3	451.5	Stadt Zürich, Geomatik + Vermessung	-	Kleiner Schlag auf bei ca. 1860mm, Band aber frei beweglich. Fehlende Schraube im Lattenkopf (links oben). Abdeckung Lattenspannfeder leicht offen (zurückgeschoben in Sollposition). Gebrauchsspuren an der ganzen Latte aber wenig Beschädigungen an Codekanten erkennbar. Mit freiem Auge nicht erkennbar, aber Latte ist scheinbar stark "konkav", Durchhangsoptimierung führt zu weit innen liegenden Auflagerpositionen, auf Grund der Position der Handgriffe ist eine weitere Verbessung nicht möglich, es verbleibt am Lattenfuss eine Überhöhung von ca. 0.6mm	0268_20201117_0925_Leica GPCL3_24514 (GeoZ 7-2)_Stadt Zürich, Geomatik + Vermessung.mat	Small damage of the code-guide at appr. 1860mm but code-tape is still free. Front, upper left screw of the staff-head is missing. Signs of wear visible throughout the whole length. Concave shape of the staff, optimization of support points not possible because of the handles in the back of the staff.
269	Leica GPCL3	24514 (GeoZ 7-2)	Stadt Zürich, Geomatik + Vermessung	FALSE	11/17/2020	14:01:39	16:22:39	141.0	55.6875	2968.6500	2	340	173	16.0	10.32	1224	1080	2100	0.000	RP	-116.56	7.108	0.389	-5.96	3.10	0.22	3.48	4.77	6.996	39.66	-115.97	0.047	-116.603	0.210	-6.21	0.026	-5.915	0.121	5.39	-115.98	0.033	-116.605	0.150	-6.21	0.019	-5.914	0.086	5.40	117640	1640	55.2	3	5.11	1234	7.9	3	4.55	2670	0.067	15	0.088	5	C	A	A	19.9	968.0	49.4	429.4	Stadt Zürich, Geomatik + Vermessung	-	Wiederholung zu 0268 wegen schlechter Lagerung. Kleiner Schlag auf bei ca. 1860mm, Band aber frei beweglich. Fehlende Schraube im Lattenkopf (links oben). Abdeckung Lattenspannfeder leicht offen (zurückgeschoben in Sollposition). Gebrauchsspuren an der ganzen Latte aber wenig Beschädigungen an Codekanten erkennbar. Mit freiem Auge nicht erkennbar, aber Latte ist scheinbar stark "konkav", Durchhangsoptimierung führt zu weit innen liegenden Auflagerpositionen, auf Grund der Position der Handgriffe ist eine weitere Verbessung nicht möglich, es verbleibt am Lattenfuss eine Überhöhung von ca. 0.6mm	0269_20201117_1348_Leica GPCL3_24514 (GeoZ 7-2)_Stadt Zürich, Geomatik + Vermessung.mat	Small damage of the code-guide at appr. 1860mm but code-tape is still free. Front, upper left screw of the staff-head is missing. Signs of wear visible throughout the whole length. Concave shape of the staff, optimization of support points not possible because of the handles in the back of the staff.
407	Leica GPCL3	24514 (GeoZ 7-2)	Stadt Zürich, Geomatik + Vermessung	FALSE	4/8/2025	9:17:19	12:11:44	174.4	55.6875	2968.6500	2	340	173	15.9	10.27	1224	920	2100	0.000	RoPr	-121.27	6.358	0.346	-6.23	3.13	0.20	3.10	5.37	6.578	35.88	-120.84	0.043	-121.295	0.188	-6.35	0.025	-6.171	0.108	5.02	-120.84	0.031	-121.296	0.134	-6.35	0.017	-6.170	0.077	5.02	117640	1740	54.1	3	4.71	1259	7.0	3	4.34	2158	0.045	15	0.045	4	C	A	A	19.7	963.9	50.6	501.9	Stadt Zürich, Geomatik + Vermessung	-	Small damage of the code-guide at appr. 1860mm but code-tape is still free. / Front, upper left screw of the staff-head is missing. / Signs of wear visible throughout the whole length. / Concave shape of the staff, optimization of support points not possible because of the handles in the back of the staff.	0407_20250408_0906_Leica GPCL3_24514 (GeoZ 7-2)_Stadt Zürich, Geomatik + Vermessung.mat	mall damage of the code-guide at appr. 1860mm but code-tape is still free. Front, upper left screw of the staff-head is missing. Signs of wear visible throughout the whole length. Staff is bent stronger than usual.
408	Leica GPCL3	24514 (GeoZ 7-2)	Stadt Zürich, Geomatik + Vermessung	FALSE	4/8/2025	14:01:32	16:55:59	174.4	55.6875	2968.6500	2	340	173	15.9	10.27	1224	900	2100	0.000	RoPr	-117.19	6.245	0.349	-6.04	3.07	0.20	3.12	5.29	6.617	36.41	-116.77	0.044	-117.203	0.192	-6.17	0.025	-5.980	0.110	5.11	-116.77	0.031	-117.222	0.142	-6.17	0.018	-5.972	0.081	5.14	117640	1665	54.1	3	4.70	1296	7.0	3	4.30	2204	0.043	15	0.049	3	X	A	A	19.8	962.4	50.7	495.5	Stadt Zürich, Geomatik + Vermessung	-	Small damage of the code-guide at appr. 1860mm but code-tape is still free. / Front, upper left screw of the staff-head is missing. / Signs of wear visible throughout the whole length. / Concave shape of the staff, optimization of support points not possible because of the handles in the back of the staff. / Repetition to 407 after trying to improve on rod shape, but as there was no success, similar support points again	0408_20250408_1348_Leica GPCL3_24514 (GeoZ 7-2)_Stadt Zürich, Geomatik + Vermessung.mat	Small damage of the code-guide at appr. 1860mm but code-tape is still free. ront, upper left screw of the staff-head is missing. gns of wear visible throughout the whole length. Staff shows stronger bend then normal.
					-117.11	±2.260			-6.21	±0.089

270	Leica GPCL2	30195 (GeoZ 7-5)	Stadt Zürich, Geomatik + Vermessung	FALSE	11/18/2020	9:12:00	10:43:06	91.1	55.6875	1934.8875	2	216	173	15.9	10.28	1224	395	1430	0.000	RP	-40.09	3.790	0.518	-1.60	2.42	0.47	3.78	19.60	4.852	28.58	-39.60	0.049	-40.212	0.400	-1.99	0.044	-1.509	0.361	4.70	-39.61	0.035	-40.212	0.283	-1.98	0.031	-1.510	0.255	4.71	74736	571	71.5	3	7.62	619	6.7	3	3.59	1524	0.052	15	0.220	1	C	B	A	19.9	967.7	49.2	484.1	Stadt Zürich, Geomatik + Vermessung	-	starke Gebrauchsspuren entlang gesamter Latte, einige kleine Schäden an Codekanten, sonst keine Auffälligkeiten	0270_20201118_0859_Leica GPVL2_30195 (GeoZ 7-5)_Stadt Zürich, Geomatik + Vermessung.mat	Signs of wear visible throughout the whole length. Minor damages to code edges in several places.
					-39.61	±0.000			-1.98	±0.000

271	GeoZ Midi (1m)	GeoZ 7-8	Stadt Zürich, Geomatik + Vermessung	FALSE	11/18/2020	12:08:28	12:53:41	45.2	45.5625	934.5375	2	107	173	15.9	10.29	1224	155	665	0.000	RP	5.89	4.742	0.576	-0.92	3.87	1.12	2.77	12.22	5.922	35.56	5.68	0.087	5.885	0.237	-0.86	0.169	-0.892	0.460	5.58	5.68	0.062	5.885	0.177	-0.87	0.120	-0.892	0.344	5.60	37022	525	70.7	3	7.64	130	6.0	3	2.63	2084	0.088	15	0.386	0	C	B	A	19.9	966.2	49.6	498.4	Stadt Zürich, Geomatik + Vermessung	-	keine besonderen Auffälligkeiten an Code oder Gehäuse	0271_20201118_1156_GeoZ Midi (1m)_GeoZ 7-8_Stadt Zürich, Geomatik + Vermessung.mat	none
273	GeoZ Midi (1m)	GeoZ 7-8	Stadt Zürich, Geomatik + Vermessung	FALSE	11/18/2020	13:48:46	14:34:34	45.8	45.5625	934.5375	2	107	173	15.9	10.30	1224	155	665	0.000	RP	4.55	3.911	0.578	-0.35	3.41	1.12	2.78	12.04	5.940	35.43	4.47	0.078	4.507	0.231	-0.23	0.152	-0.277	0.449	5.02	4.48	0.056	4.507	0.177	-0.24	0.108	-0.277	0.344	5.05	37022	524	70.6	3	7.61	134	6.0	3	2.63	2096	0.255	15	0.782	0	C	B	A	20.0	965.5	49.7	509.4	Stadt Zürich, Geomatik + Vermessung	-	Wiederholung zu 271 (Hin- und Retour zeigte relativ grossen Unterschied). Keine besonderen Auffälligkeiten an Code oder Gehäuse	0273_20201118_1334_GeoZ Midi (1m)_GeoZ 7-8_Stadt Zürich, Geomatik + Vermessung.mat	none
432	GeoZ Midi (0.9m)	GeoZ 7-8	Stadt Zürich, Geomatik + Vermessung	FALSE	4/16/2025	9:04:40	9:59:09	54.5	45.5625	934.5375	2	107	173	15.9	10.27	1224	160	660	0.000	RoPr	-9.45	4.119	0.613	-0.33	4.10	1.19	2.93	10.70	6.384	36.90	-9.53	0.087	-9.508	0.210	-0.25	0.170	-0.273	0.412	5.55	-9.54	0.062	-9.508	0.160	-0.24	0.120	-0.273	0.316	5.56	37022	581	71.5	3	7.64	468	5.6	3	2.51	2070	0.034	15	0.073	0	C	B	A	19.6	945.6	51.1	461.1	Stadt Zürich, Geomatik + Vermessung	-	None.	0432_20250416_0853_GeoZ Midi (0.9m)_GeoZ 7-8_Stadt Zürich, Geomatik + Vermessung.mat	None.
433	GeoZ Midi (0.9m)	GeoZ 7-8	Stadt Zürich, Geomatik + Vermessung	FALSE	4/16/2025	10:45:26	11:39:55	54.5	45.5625	934.5375	2	107	173	15.9	10.27	1224	160	660	0.000	RoPr	-7.02	4.038	0.618	0.21	4.16	1.20	2.96	10.64	6.400	36.99	-7.08	0.087	-7.031	0.219	0.25	0.169	0.181	0.430	5.53	-7.08	0.061	-7.031	0.160	0.25	0.120	0.181	0.314	5.53	37022	529	69.9	3	7.60	466	5.4	3	2.42	2085	0.029	15	0.041	0	C	B	A	19.6	946.8	51.1	460.3	Stadt Zürich, Geomatik + Vermessung	-	None.	0433_20250416_1033_GeoZ Midi (0.9m)_GeoZ 7-8_Stadt Zürich, Geomatik + Vermessung.mat	None.
					-1.61	±6.765			-0.27	±0.400

274	GeoZ Mini (0.2m)	GeoZ 7-7	Stadt Zürich, Geomatik + Vermessung	FALSE	11/18/2020	15:49:48	16:00:33	10.8	45.5625	215.6625	2	26	173	15.9	10.30	1224	25	225	0.000	RP	25.08	3.351	1.118	-13.95	23.32	8.03	1.92	13.58	5.075	32.74	24.99	0.149	25.054	0.295	-13.86	1.069	-13.808	2.121	3.51	25.00	0.105	25.054	0.209	-13.90	0.756	-13.808	1.503	3.51	8996	35	60.2	3	4.73	58	7.5	3	3.22	322	0.235	15	0.635	0	X	X	X	20.0	965.1	49.6	562.7	Stadt Zürich, Geomatik + Vermessung	-	keine Auffälligkeiten	0274_20201118_1535_GeoZ Mini (0.2m)_GeoZ 7-7_Stadt Zürich, Geomatik + Vermessung.mat	none
275	GeoZ Mini (0.2m)	GeoZ 7-7	Stadt Zürich, Geomatik + Vermessung	FALSE	11/18/2020	16:36:03	16:47:12	11.1	45.5625	215.6625	2	26	173	15.9	10.30	1224	25	225	0.000	RP	24.73	3.501	1.230	-13.16	16.69	8.84	2.12	13.23	5.149	31.63	24.51	0.155	25.023	0.540	-13.32	1.117	-14.792	3.856	3.65	24.52	0.110	25.045	0.383	-13.32	0.794	-14.916	2.736	3.67	8996	35	58.0	3	4.59	59	7.2	3	3.18	291	0.596	15	1.366	0	X	X	X	20.0	965.0	49.6	578.2	Stadt Zürich, Geomatik + Vermessung	-	Wiederholung zu 0274. keine Auffälligkeiten an der Latte	0275_20201118_1623_GeoZ Mini (0.2m)_GeoZ 7-7_Stadt Zürich, Geomatik + Vermessung.mat	none
412	GeoZ Mini	GeoZ 7-7	Stadt Zürich, Geomatik + Vermessung	FALSE	4/9/2025	14:23:18	14:36:02	12.7	45.5625	215.6625	2	26	173	15.9	10.26	1224	25	225	0.000	RoPr	11.06	3.274	1.016	-15.74	14.94	7.28	1.73	13.78	4.917	32.51	11.14	0.151	11.057	0.164	-17.16	1.087	-15.793	1.177	3.52	11.13	0.107	11.057	0.167	-17.09	0.771	-15.793	1.201	3.53	8996	48	57.6	3	4.56	80	6.9	3	3.21	385	0.032	15	0.023	0	X	X	X	19.8	959.9	50.8	498.7	Stadt Zürich, Geomatik + Vermessung	-	Spirit level only available on one side.	0412_20250409_1411_GeoZ Mini_GeoZ 7-7_Stadt Zürich, Geomatik + Vermessung.mat	Spirit level only available on one side.
413	GeoZ Mini	GeoZ 7-7	Stadt Zürich, Geomatik + Vermessung	FALSE	4/9/2025	15:24:19	15:36:56	12.6	45.5625	215.6625	2	26	173	15.9	10.26	1224	25	225	0.000	RoPr	14.92	3.441	1.044	-19.66	23.18	7.47	1.77	13.82	4.873	31.87	14.99	0.170	14.900	0.157	-21.25	1.220	-19.806	1.130	3.99	14.99	0.120	14.900	0.122	-21.26	0.862	-19.806	0.875	3.99	8996	53	57.7	3	4.59	69	6.9	3	3.22	361	0.035	15	0.053	0	X	X	X	19.8	959.5	50.6	526.8	Stadt Zürich, Geomatik + Vermessung	-	Spirit level only available on one side. / Staff about 20px off center (1244 instead of 1224) hopefulla avoids settling isse due to heady-heavy construction.	0413_20250409_1513_GeoZ Mini_GeoZ 7-7_Stadt Zürich, Geomatik + Vermessung.mat	Spirit level only available on one side.
414	GeoZ mini (0.2m)	GeoZ 7-7	Stadt Zürich, Geomatik + Vermessung	FALSE	4/10/2025	11:01:10	11:13:51	12.7	45.5625	215.6625	2	26	173	15.9	10.26	1224	20	220	0.000	RoPr	16.33	3.091	1.003	-20.17	19.91	7.19	1.72	13.87	4.847	30.67	16.33	0.152	16.332	0.150	-21.23	1.093	-20.269	1.082	3.59	16.34	0.108	16.332	0.109	-21.27	0.773	-20.269	0.783	3.58	8996	64	56.8	3	4.26	103	6.7	3	3.21	344	0.049	15	0.026	0	X	X	X	19.7	964.4	50.7	482.6	Stadt Zürich, Geomatik + Vermessung	-	Testrun with rubberized support points to avoid slippage of the staff during measurement	0414_20250410_1048_GeoZ mini (0.2m)_GeoZ 7-7_Stadt Zürich, Geomatik + Vermessung.mat	none
415	GeoZ mini (0.2m)	GeoZ 7-7	Stadt Zürich, Geomatik + Vermessung	FALSE	4/10/2025	11:46:37	11:59:19	12.7	45.5625	215.6625	2	26	173	15.9	10.26	1224	20	220	0.000	RoPr	14.58	3.111	1.009	-16.78	20.06	7.24	1.73	13.92	4.838	30.71	14.55	0.151	14.560	0.154	-17.66	1.087	-16.695	1.109	3.56	14.55	0.107	14.560	0.121	-17.61	0.770	-16.695	0.868	3.57	8996	62	56.8	3	4.26	102	6.7	3	3.23	336	0.056	15	0.052	0	X	X	X	19.7	964.4	50.7	483.0	Stadt Zürich, Geomatik + Vermessung	-	Testrun with rubberized support points to avoid slippage of the staff during measurement. / Rubberized Support seems to work. Repeat to be sure.	0415_20250410_1135_GeoZ mini (0.2m)_GeoZ 7-7_Stadt Zürich, Geomatik + Vermessung.mat	Bubble level only available on one side of the staff levelller.
419	GeoZ mini (0.2m)	GeoZ 7-7	Stadt Zürich, Geomatik + Vermessung	FALSE	4/11/2025	11:27:20	11:40:12	12.9	45.5625	215.6625	2	26	173	15.9	10.26	1224	20	220	0.000	RoPr	14.41	3.103	0.995	-14.17	20.14	7.14	1.71	13.95	4.776	29.22	14.40	0.153	14.395	0.141	-15.05	1.099	-14.046	1.011	3.61	14.40	0.108	14.395	0.104	-15.08	0.776	-14.046	0.747	3.60	8996	65	56.8	3	4.25	101	6.7	3	3.22	322	0.042	15	0.025	0	X	X	X	19.7	961.3	50.6	506.7	Stadt Zürich, Geomatik + Vermessung	-	First Testrun with CodeGuideCenter-Monitoring.	0419_20250411_1116_GeoZ mini (0.2m)_GeoZ 7-7_Stadt Zürich, Geomatik + Vermessung.mat	none
421	GeoZ Mini (0.2m)	GeoZ 7-7	Stadt Zürich, Geomatik + Vermessung	FALSE	4/11/2025	14:54:09	15:06:54	12.7	45.5625	215.6625	2	26	173	15.9	10.26	1224	20	220	0.000	RoPr	11.26	3.095	1.011	-13.27	20.10	7.24	1.73	13.96	4.839	30.73	11.21	0.153	11.231	0.157	-14.05	1.095	-13.063	1.130	3.59	11.22	0.108	11.251	0.109	-14.16	0.775	-13.045	0.786	3.59	8996	63	56.8	3	4.27	97	6.7	3	3.22	335	0.057	15	0.119	0	X	X	X	19.7	959.0	50.8	499.1	Stadt Zürich, Geomatik + Vermessung	-	2nd Testrun for Code Guide Center Monitoring	0421_20250411_1442_GeoZ Mini (0.2m)_GeoZ 7-7_Stadt Zürich, Geomatik + Vermessung.mat	2nd Testrun for Code Guide Center Monitoring.
					16.52	±5.047			-16.71	±2.978

277	Leica GPCL 3	24508 (GeoZ 6-3)	Stadt Zürich, Geomatik + Vermessung	FALSE	11/24/2020	11:25:48	13:57:54	152.1	55.6875	2968.6500	2	340	173	15.9	10.29	1224	650	2600	0.000	RP	-95.28	6.009	0.353	-3.89	2.29	0.20	3.18	11.63	5.290	28.67	-95.17	0.041	-95.296	0.239	-3.98	0.023	-3.877	0.137	4.78	-95.17	0.029	-95.315	0.176	-3.98	0.017	-3.872	0.101	4.84	117640	1000	59.9	3	5.55	967	6.4	3	3.70	3119	0.583	15	1.104	0	C	B	A	20.1	963.7	49.3	461.1	Stadt Zürich, Geomatik + Vermessung	-	Wiederholung zu 276 (abgebrochen wegen Interferometerfehler), Auskrichtung kontrolliert aber nicht verändert. Schraube vorne links oben im Lattenkopf fehlt. Starke Gebrauchsspuren, aber optisch keine groben Schäden am £Gehäuse. codeband frei gängig. Kratzer im Code bei ca. 1650, 1825 & 2000mm.	0277_20201124_1112_Leica GPCL 3_24508 (GeoZ 6-3)_Stadt Zürich, Geomatik + Vermessung.mat	Strong signs of use, but no major defects visible. Screw on the staff-head missing (top, left).
390	Leica GPCL3	24508 (GeoZ 6-3)	Stadt Zürich, Geomatik + Vermessung	FALSE	3/24/2025	9:00:20	11:56:26	176.1	55.6875	2967.6375	2	340	173	15.9	10.26	1224	670	2500	0.000	RoPr	-97.52	2.532	0.331	-2.67	2.52	0.19	2.97	11.26	5.231	27.85	-97.51	0.042	-97.581	0.217	-2.70	0.024	-2.667	0.125	5.02	-97.51	0.030	-97.581	0.154	-2.70	0.017	-2.667	0.089	5.03	117640	1507	64.0	3	6.21	1250	7.3	3	4.48	2852	0.100	15	0.066	0	C	B	A	19.8	952.9	50.6	452.2	Stadt Zürich, Geomatik + Vermessung	-	Heavy signs of wear (hard to find "clean edge" for camera calibration). / A lot of small scratches all over the code (mostly vertical). / Cacked, hard to remove dirt on the fist 50mm or so. / Minor defect in staff housing near handle (left side). / Cover for tensioning spring dislocated. / Screw for staff tripod loose (nut not fixed in housing anymore) and Housing-screw on the top left missing.	0390_20250324_0848_Leica GPCL3_24508 (GeoZ 6-3)_Stadt Zürich, Geomatik + Vermessung.mat	Heavy signs of wear. A lot of small scratches all over the code (mostly vertical). Cacked, hard to remove dirt on the fist 50mm or so. Minor defect in staff housing near handle (left side). Cover for tensioning spring dislocated. Screw for staff tripod loose (nut not fixed in housing anymore) and Housing-screw on the top left missing.
					-96.34	±1.166			-3.34	±0.640

279	Leica GPCL 3	24512 (GeoZ 6-4)	Stadt Zürich, Geomatik + Vermessung	FALSE	11/25/2020	8:46:19	11:14:26	148.1	55.6875	2968.6500	2	340	173	15.9	10.30	1224	700	2600	0.000	RP	-132.54	3.340	0.348	-3.25	0.86	0.20	3.07	10.97	5.223	30.07	-132.43	0.035	-132.583	0.211	-3.28	0.020	-3.223	0.121	4.02	-132.43	0.025	-132.585	0.151	-3.28	0.014	-3.223	0.087	4.04	117640	1448	61.5	3	5.45	1986	6.6	3	3.74	4788	0.296	15	0.739	0	C	B	A	20.0	961.0	49.2	445.5	Stadt Zürich, Geomatik + Vermessung	-	Wiederholung zu 278 (abgebrochen mit unbekanntem Interferometerfehler). Starke Gebrauchsspuren, aber keine groben Beschädigungen mit freiem Aug erkennbar. Leichte "rote" Kratzer am Code bei ca. 1890 bis 1970mm.	0279_20201125_0830_Leica GPCL 3_24512 (GeoZ 6-4)_Stadt Zürich, Geomatik + Vermessung.mat	Strong signs of use, but no major defects visible. Small red scratches/markings from appr. 1890 to 1970mm on the code.
391	Leica GPCL3	24512 (GeoZ 6-4)	Stadt Zürich, Geomatik + Vermessung	FALSE	3/24/2025	16:21:11	19:16:34	175.4	55.6875	2968.6500	2	340	173	15.9	10.26	1224	700	2600	0.000	RoPr	-141.56	2.302	0.340	-3.14	0.99	0.19	2.99	10.80	5.058	28.62	-141.59	0.037	-141.700	0.210	-3.06	0.021	-3.106	0.120	4.28	-141.59	0.026	-141.700	0.149	-3.06	0.015	-3.106	0.085	4.29	117640	1417	62.7	3	5.50	2323	6.3	3	3.70	4468	0.083	15	0.066	0	C	B	A	19.8	953.9	50.6	457.0	Stadt Zürich, Geomatik + Vermessung	-	Heavy signs of wear. / Top right housing screw was loose.	0391_20250324_1608_Leica GPCL3_24512 (GeoZ 6-4)_Stadt Zürich, Geomatik + Vermessung.mat	Heavy signs of wear. Top right housing screw was loose.
					-137.01	±4.579			-3.17	±0.110

280	Leica GPCL 2	76183	Stadt Zürich, Geomatik + Vermessung	FALSE	11/25/2020	15:02:16	16:35:04	92.8	55.6875	1934.8875	2	216	173	15.9	10.30	1224	445	1555	0.000	RP	5.60	6.879	0.433	3.14	5.53	0.39	3.16	-2.28	8.452	37.07	5.56	0.061	5.550	0.285	3.19	0.055	3.189	0.255	5.89	5.57	0.043	5.550	0.202	3.18	0.039	3.189	0.181	5.90	74736	251	67.7	3	5.47	271	7.5	3	3.07	2407	0.107	15	0.412	0	C	B	A	20.0	960.3	49.2	453.4	Stadt Zürich, Geomatik + Vermessung	-	Keine auffälligen Schäden an Code oder Gehäuse. Halteband für Kalibrierung entfernt.	0280_20201125_1450_Leica GPCL 2_76183_Stadt Zürich, Geomatik + Vermessung.mat	none
394	Leica GPCL2	76183	Stadt Zürich, Geomatik + Vermessung	FALSE	3/28/2025	9:55:21	11:47:22	112.0	55.6875	1934.8875	2	216	173	15.9	10.27	1224	450	1555	0.000	RoPr	3.98	5.606	0.433	3.74	4.13	0.39	3.13	-5.88	8.443	37.63	3.95	0.056	3.907	0.268	3.78	0.050	3.796	0.240	5.34	3.95	0.040	3.907	0.189	3.78	0.036	3.796	0.170	5.35	74736	730	70.5	3	6.33	1115	7.3	3	4.45	2146	0.034	15	0.033	0	C	A	A	19.9	951.9	50.7	466.1	Stadt Zürich, Geomatik + Vermessung	-	Very dirty @ 125 to 145cm. Whole code cleaned. / Strong scratches @ 74, 131.7, 150, 155 & 177.5cm. / Paint scratched away from guide groove @ 83.5cm (Front/Right).	0394_20250328_0941_Leica GPCL2_76183_Stadt Zürich, Geomatik + Vermessung.mat	Very dirty @ 125 to 145cm. Whole code cleaned. Strong scratches @ 74, 131.7, 150, 155 & 177.5cm. Paint scratched away from guide groove @ 83.5cm (Front/Right).
					4.76	±0.807			3.48	±0.298

281	GeoZ Midi	GeoZ 2-9	Stadt Zürich, Geomatik + Vermessung	FALSE	11/30/2020	11:05:09	11:50:33	45.4	55.6875	934.5375	2	106	173	15.9	10.29	1224	180	700	0.000	RP	23.31	3.521	0.445	-3.31	7.51	0.86	2.14	9.13	4.920	26.78	23.46	0.065	23.381	0.216	-3.95	0.127	-3.465	0.418	4.13	23.46	0.046	23.381	0.159	-3.95	0.090	-3.465	0.307	4.15	36676	430	68.8	3	6.53	113	7.4	3	2.97	1523	0.054	15	0.245	0	C	A	A	20.1	964.8	48.7	473.0	Stadt Zürich, Geomatik + Vermessung	-	keine	0281_20201130_1053_GeoZ Midi_GeoZ 2-9_Stadt Zürich, Geomatik + Vermessung.mat	none
282	GeoZ Midi	GeoZ 2-9	Stadt Zürich, Geomatik + Vermessung	FALSE	11/30/2020	14:57:09	15:42:20	45.2	45.5625	934.5375	2	107	173	15.9	10.30	1224	180	700	0.000	RP	37.27	3.908	0.437	-3.03	8.67	0.85	2.12	9.23	4.834	26.63	37.46	0.062	37.390	0.218	-3.80	0.121	-3.276	0.424	4.00	37.44	0.045	37.394	0.186	-3.71	0.088	-3.278	0.362	4.09	37022	371	69.5	3	6.25	133	7.1	3	2.82	1620	0.025	15	0.087	1	C	A	A	20.2	963.6	48.2	494.1	Stadt Zürich, Geomatik + Vermessung	-	Wiederholung zu 0281, auf Grund relativ grosser Differenz Vor und Rückmessung. Ausrichtung kontrolliert und licht angepasst. keine	0282_20201130_1445_GeoZ Midi_GeoZ 2-9_Stadt Zürich, Geomatik + Vermessung.mat	none
283	GeoZ Midi	GeoZ 2-9	Stadt Zürich, Geomatik + Vermessung	FALSE	12/1/2020	9:27:02	10:12:28	45.4	45.5625	934.5375	2	107	173	15.9	10.29	1224	180	700	0.000	RP	34.09	3.687	0.439	-2.51	7.96	0.85	2.12	9.08	4.912	27.01	34.31	0.064	34.204	0.206	-3.34	0.125	-2.754	0.401	4.09	34.31	0.046	34.204	0.174	-3.32	0.090	-2.754	0.339	4.16	37022	425	65.0	3	6.28	104	6.9	3	2.75	1866	0.040	15	0.204	0	C	B	A	20.1	955.4	48.7	443.8	Stadt Zürich, Geomatik + Vermessung	-	Testmessung zur Kontrolle der Lattenbewegung auf Grund Vibrationen	0283_20201201_0915_GeoZ Midi_GeoZ 2-9_Stadt Zürich, Geomatik + Vermessung.mat	none
409	GeoZ Midi	GeoZ 2-9	Stadt Zürich, Geomatik + Vermessung	FALSE	4/9/2025	9:20:44	10:14:29	53.8	55.6875	918.3375	2	105	173	15.9	10.27	1224	145	700	0.000	RoPr	32.55	3.258	0.449	-2.07	7.24	0.87	2.13	6.80	4.932	26.87	32.68	0.063	32.611	0.214	-2.60	0.123	-2.167	0.418	3.97	32.66	0.045	32.598	0.164	-2.56	0.088	-2.152	0.319	3.99	36330	346	66.9	3	6.20	108	6.5	3	2.68	1662	0.014	15	0.011	3	C	A	A	19.7	960.4	50.7	487.2	Stadt Zürich, Geomatik + Vermessung	-	none	0409_20250409_0907_GeoZ Midi_GeoZ 2-9_Stadt Zürich, Geomatik + Vermessung.mat	none
410	GeoZ Midi	GeoZ 2-9	Stadt Zürich, Geomatik + Vermessung	FALSE	4/9/2025	11:00:46	11:55:13	54.5	45.5625	934.5375	2	107	173	15.9	10.27	1224	145	700	0.000	RoPr	31.18	3.461	0.437	-1.99	8.36	0.85	2.14	6.88	4.977	27.25	31.36	0.063	31.243	0.199	-2.65	0.123	-2.124	0.388	4.06	31.36	0.044	31.243	0.141	-2.65	0.087	-2.124	0.275	4.06	37022	374	66.7	3	6.16	127	6.6	3	2.73	1573	0.020	15	0.015	0	C	A	A	19.8	960.6	50.8	487.5	Stadt Zürich, Geomatik + Vermessung	-	none	0410_20250409_1045_GeoZ Midi_GeoZ 2-9_Stadt Zürich, Geomatik + Vermessung.mat	none
					31.85	±4.659			-3.24	±0.556

285	GeoZ Mini	GeoZ 2-6	Stadt Zürich, Geomatik + Vermessung	FALSE	12/1/2020	13:43:55	13:54:24	10.5	45.5625	215.6625	2	26	173	15.9	10.30	1224	20	225	0.000	RP	33.21	3.576	0.989	-26.12	19.72	7.11	1.72	-0.71	4.043	25.29	32.96	0.131	33.098	0.312	-24.70	0.943	-25.416	2.244	3.10	32.96	0.092	33.098	0.217	-24.69	0.667	-25.416	1.558	3.10	8996	40	67.9	3	6.42	28	7.8	3	2.95	303	0.016	15	0.017	0	X	X	X	20.2	955.0	48.8	453.3	Stadt Zürich, Geomatik + Vermessung	-	none	0285_20201201_1332_GeoZ Mini_GeoZ 2-6_Stadt Zürich, Geomatik + Vermessung.mat	none
286	GeoZ Mini	GeoZ 2-6	Stadt Zürich, Geomatik + Vermessung	FALSE	12/1/2020	14:49:25	14:59:59	10.6	45.5625	215.6625	2	26	173	15.9	10.30	1224	20	225	0.000	RP	30.77	4.116	0.985	-23.18	19.23	7.09	1.71	-0.69	4.034	22.34	30.46	0.140	30.660	0.308	-21.50	1.007	-22.559	2.216	3.32	30.46	0.099	30.660	0.221	-21.50	0.714	-22.559	1.593	3.33	8996	41	67.9	3	6.33	30	7.8	3	2.97	295	0.042	15	0.199	0	X	X	X	20.1	955.2	48.7	459.6	Stadt Zürich, Geomatik + Vermessung	-	Wiederholung zu 0285 (reine Kontrolle)	0286_20201201_1437_GeoZ Mini_GeoZ 2-6_Stadt Zürich, Geomatik + Vermessung.mat	none
424	GeoZ Mini	GeoZ 2-6	Stadt Zürich, Geomatik + Vermessung	FALSE	4/14/2025	14:41:47	14:54:35	12.8	45.5625	215.6625	2	26	173	15.9	10.27	1224	20	220	0.000	RoPr	27.06	2.588	0.964	-13.41	17.82	6.93	1.68	-1.32	3.862	22.46	26.91	0.127	27.053	0.295	-12.80	0.914	-13.469	2.125	3.01	26.89	0.089	27.221	0.196	-12.72	0.644	-14.492	1.397	3.00	8996	42	67.5	3	5.89	45	7.3	3	2.89	278	0.037	15	0.021	0	X	X	X	19.6	948.2	50.9	459.4	Stadt Zürich, Geomatik + Vermessung	-	none	0424_20250414_1430_GeoZ Mini_GeoZ 6-2_Stadt Zürich, Geomatik + Vermessung.mat	none
425	GeoZ Mini	GeoZ 2-6	Stadt Zürich, Geomatik + Vermessung	FALSE	4/14/2025	15:31:49	15:44:31	12.7	45.5625	215.6625	2	26	173	15.9	10.27	1224	20	220	0.000	RoPr	31.17	2.545	0.964	-21.61	17.20	6.93	1.67	-1.39	3.866	22.66	31.03	0.125	31.175	0.294	-21.11	0.901	-21.719	2.114	2.96	31.03	0.089	31.348	0.203	-21.19	0.639	-22.769	1.451	2.98	8996	45	67.5	3	5.88	48	7.3	3	2.89	286	0.037	15	0.025	0	X	X	X	19.6	947.4	51.1	460.0	Stadt Zürich, Geomatik + Vermessung	-	none	0425_20250414_1520_GeoZ Mini_GeoZ 6-2_Stadt Zürich, Geomatik + Vermessung.mat	none
					30.34	±2.193			-20.03	±4.432

295	Trimble LD13	61530 (FHNW No. 61)	Fachhochschule Nordwestschweiz	FALSE	4/28/2021	10:11:21	12:08:07	116.8	52.2500	2965.0000	2	266	171	15.8	10.34	1224	420	2700	0.000	RP	-12.67	4.734	0.487	1.03	2.84	0.28	3.58	-15.12	8.583	49.43	-12.48	0.067	-12.591	0.166	0.83	0.039	0.995	0.097	6.49	-12.48	0.048	-12.581	0.122	0.83	0.028	0.999	0.071	6.50	90972	2369	64.4	3	7.81	255	9.9	3	4.20	7641	0.119	15	0.152	0	X	B	A	20.3	945.7	48.4	443.8	Fachhochschule Nordwestschweiz	-	keine sichtbaren Schäden, Lagerpunkte relativ weit aussen	0295_20210428_0958_Trimble LD13_61530 (FHNW No. 61)_Fachhochschule Nordwestschweiz.mat	none
					-12.48	±0.000			0.83	±0.000

298	Trimble LD13	61534 (FHNW No. 62)	Fachhochschule Nordwestschweiz	FALSE	4/29/2021	13:23:03	15:21:34	118.5	52.2500	2965.0000	2	266	173	15.9	10.30	1224	600	2450	0.000	RP	2.12	5.743	0.602	1.58	3.39	0.36	4.54	-15.39	8.737	49.98	2.10	0.069	2.012	0.235	1.64	0.041	1.663	0.136	6.83	2.11	0.049	2.015	0.168	1.64	0.029	1.662	0.098	6.84	92036	1269	65.0	3	7.91	204	9.9	3	4.21	6902	0.140	15	0.281	1	X	B	A	19.9	948.1	49.0	453.1	Fachhochschule Nordwestschweiz	-	keine sichtbaren Schäden. Wiederholung zu 0296 und 0297, beide wegen zu grosser "Schere". Ausrichtung kontrolliert, aber nicht verändert.	0298_20210429_1309_Trimble LD13_61534 (FHNW No. 62)_Fachhochschule Nordwestschweiz.mat	none
					2.11	±0.000			1.64	±0.000

300	Trimble LD13	61528 (FHNW No. 52)	Fachhochschule Nordwestschweiz	FALSE	4/30/2021	13:28:51	15:29:05	120.2	52.2500	2965.0000	2	266	173	15.9	10.30	1224	530	2550	0.000	RP	-2.52	4.929	0.505	2.54	1.94	0.29	3.80	-11.93	8.458	47.52	-2.49	0.068	-2.512	0.178	2.57	0.040	2.561	0.103	6.78	-2.49	0.048	-2.516	0.129	2.57	0.028	2.563	0.075	6.78	92036	1223	64.6	3	6.88	601	9.8	3	3.91	5375	0.115	15	0.271	1	C	B	A	19.9	950.0	48.9	441.5	Fachhochschule Nordwestschweiz	-	optisch keine Besonderheiten	0300_20210430_1317_Trimble LD13_61528 (FHNW No. 52)_Fachhochschule Nordwestschweiz.mat	none
					-2.49	±0.000			2.57	±0.000

302	Trimble LD13	61527 (FHNW No. 51)	Fachhochschule Nordwestschweiz	FALSE	5/3/2021	13:40:58	15:41:13	120.3	52.2500	2965.0000	2	266	173	16.0	10.30	1224	400	2750	0.000	RP	-21.63	4.862	0.515	0.84	1.27	0.30	3.94	-11.98	8.296	47.62	-21.87	0.064	-21.650	0.197	0.85	0.037	0.864	0.115	6.41	-21.87	0.045	-21.651	0.143	0.85	0.026	0.864	0.083	6.42	92036	1259	64.4	3	6.80	503	10.4	3	5.26	3397	0.140	15	0.283	0	C	B	A	20.1	960.6	48.6	423.1	Fachhochschule Nordwestschweiz	-	Wiederholung zu 0301 wegen zu grosser Schere.	0302_20210503_1329_Trimble LD13_61527 (FHNW No. 51)_Fachhochschule Nordwestschweiz.mat	none
					-21.87	±0.000			0.85	±0.000

305	Leica GPCL2	79716	Sersa Group AG, BIM Reality Capture	FALSE	2/25/2022	9:39:42	11:13:04	93.4	55.6875	1934.8875	2	216	173	15.9	10.30	1224	446	1554	0.000	RP	-4.81	4.190	0.382	3.48	2.47	0.34	2.72	-14.36	4.107	24.04	-5.09	0.051	-4.898	0.182	3.41	0.046	3.559	0.164	4.81	-5.09	0.036	-4.898	0.130	3.41	0.032	3.558	0.117	4.81	74736	635	61.0	3	8.13	52	9.5	3	4.07	3999	0.106	15	0.377	0	C	B	A	20.0	965.9	48.6	466.4	Sersa Group AG, BIM Reality Capture	-	Staff very dirty upon delivery. Code cleaned.	0305_20220225_0926_Leica GPCL2_79716_Sersa Group AG, BIM Reality Capture.mat	none
					-5.09	±0.000			3.41	±0.000

306	Leica GPCL2	79720	Sersa Group AG, BIM Reality Capture	FALSE	2/25/2022	12:25:09	13:58:36	93.4	55.6875	1934.8875	2	216	173	15.9	10.30	1224	370	1630	0.000	RP	-0.62	2.687	0.279	4.38	0.30	0.25	2.03	-4.40	3.264	18.15	-0.59	0.034	-0.635	0.135	4.35	0.030	4.385	0.122	3.24	-0.58	0.024	-0.635	0.100	4.34	0.022	4.385	0.089	3.25	74736	542	66.0	3	8.02	73	9.0	3	3.63	4253	0.064	15	0.253	0	C	A	A	20.1	965.4	48.2	443.4	Sersa Group AG, BIM Reality Capture	-		0306_20220225_1213_Leica GPCL2_79720_Sersa Group AG, BIM Reality Capture.mat	none
					-0.58	±0.000			4.34	±0.000

320	WILD GPCL3	9685	Basel Landschaft, Amt für Geoinformation	FALSE	3/15/2022	8:54:57	11:23:28	148.5	55.6875	2968.6500	2	340	173	15.9	10.30	1224	600	2400	0.000	RP	16.16	3.080	0.233	3.00	1.86	0.13	2.08	8.66	5.114	34.23	16.32	0.031	16.112	0.155	3.02	0.018	3.020	0.089	3.60	16.33	0.022	16.112	0.121	3.01	0.013	3.020	0.069	3.67	117640	705	62.6	3	3.22	2049	7.6	3	4.17	997	0.101	15	0.178	0	C	B	A	20.0	965.8	48.7	451.4	Basel Landschaft, Amt für Geoinformation	-	Invarband eher schwergängig; starke Gebrauchsspuren; leichte Gehäusedeformation aussen links bei ~220cm; Fussplatte verschweisst?; sehr viele Schraubencuvetten gebrochen	0320_20220315_0842_WILD GPCL3_9685_Basel Landschaft, Amt für Geoinformation.mat	Strong signs of use Invar-Codetape does not move very easily
325	WILD GPCL3	9685	Basel Landschaft, Amt für Geoinformation	FALSE	3/16/2022	8:47:34	11:15:11	147.6	55.6875	2968.6500	2	340	173	15.9	10.30	1224	600	2400	0.000	RP	18.30	3.066	0.216	2.38	1.67	0.12	1.93	8.70	5.129	34.29	18.48	0.030	18.291	0.118	2.38	0.017	2.382	0.068	3.49	18.48	0.022	18.285	0.104	2.37	0.012	2.381	0.059	3.60	117640	863	62.7	3	3.50	1611	7.7	3	4.32	812	0.083	15	0.084	2	C	B	A	20.2	963.2	48.8	489.7	Basel Landschaft, Amt für Geoinformation	-		0325_20220316_0836_WILD GPCL3_9685_Basel Landschaft, Amt für Geoinformation.mat	Invar-Tape seems blocked within housing (does not move freely, creates unfamilar sounds when moving the staff). Replacment/Repair of staff recommended for precision levelling.
					17.40	±1.074			2.69	±0.320

324	WILD GPCL3	9981	Basel Landschaft, Amt für Geoinformation	FALSE	3/15/2022	16:22:42	18:48:09	145.4	55.6875	2968.6500	2	340	173	15.9	10.30	1224	500	2340	0.000	RP	-23.91	4.336	0.344	6.58	1.74	0.20	3.07	20.40	8.360	50.51	-23.18	0.036	-23.896	0.161	6.26	0.021	6.577	0.092	4.19	-23.17	0.026	-23.894	0.114	6.26	0.015	6.577	0.065	4.19	117640	2680	60.2	3	5.21	1822	8.4	3	5.04	1852	0.057	15	0.051	3	X	C	A	20.1	961.9	49.0	486.8	Basel Landschaft, Amt für Geoinformation	-	starke Codeschäden <50cm, 14 & 225cm; Lattengehäuse links aussen leicht deformiert bei ca. 110cm; Kamerakalibrierung ev. problematisch (Knick in allen Kanten); Lattengehäuse nahe Fuss scheint stark gekrümmt.	0324_20220315_1606_WILD GPCL3_9981_Basel Landschaft, Amt für Geoinformation.mat	Strong signs of wear on the digital code (yellow paint heavily chipped and scratched at several positions). Replacement/Repair of staff recommended for precision levelling.
					-23.17	±0.000			6.26	±0.000

382	Leica GPCL3	80912 (A)	Basel Landschaft, Amt für Geoinformation	FALSE	3/5/2025	11:48:32	14:44:35	176.1	55.6875	2968.6500	2	340	173	15.9	10.27	1224	430	2400	0.000	RoPr	-13.46	3.433	0.286	1.42	0.89	0.16	2.53	-25.17	3.548	19.97	-13.42	0.031	-13.499	0.164	1.40	0.018	1.421	0.094	3.61	-13.43	0.022	-13.511	0.118	1.41	0.013	1.430	0.068	3.62	117640	926	63.0	3	5.98	425	7.4	3	3.32	6021	0.047	15	0.056	13	C	C	A	20.1	964.0	50.2	504.4	Basel Landschaft, Amt für Geoinformation	-	Intensive rust-stains in the rear part of the 0-point excenter. / Strong dent in the rear of the housing (left side) @ ca. 423mm. / Cover for tensioning screw missing on top of the staff. / 381 interrupted by motor communications error.	0382_20250305_1137_Leica GPCL3_80912 (A)_Basel Landschaft, Amt für Geoinformation.mat	Intensive rust-stains in the rear part of the 0-point excenter. Strong dent in the rear of the housing (left side) @ ca. 423mm. Cover for tensioning screw missing on top of the staff.
384	Leica GPCL3	80912 (A)	Basel Landschaft, Amt für Geoinformation	FALSE	3/7/2025	9:15:45	12:11:28	175.7	55.6875	2968.6500	2	340	173	15.9	10.27	1224	380	2230	0.000	RoPr	-7.08	3.212	0.284	1.12	1.01	0.16	2.55	-25.15	3.556	20.21	-7.15	0.029	-7.098	0.160	1.14	0.017	1.108	0.092	3.48	-7.15	0.021	-7.099	0.117	1.14	0.012	1.109	0.067	3.50	117640	918	63.0	3	6.09	399	7.4	3	3.32	6145	0.040	15	0.036	1	C	C	A	20.1	959.3	50.3	493.6	Basel Landschaft, Amt für Geoinformation	-	Rust stains on the back of the 0-point excenter. / Strong dent in the rear of the housing (left side) @ ca. 423mm. / Cover for tensioning screw missing on top of the staff.	0384_20250307_0904_Leica GPCL3_80912 (A)_Basel Landschaft, Amt für Geoinformation.mat	Rust stains on the back of the 0-point excenter. Strong dent in the rear of the housing (left side) @ ca. 423mm. Cover for tensioning screw missing on top of the staff.
387	Leica GPCL3	80912 (A)	Basel Landschaft, Amt für Geoinformation	FALSE	3/12/2025	10:01:39	12:58:00	176.3	55.6875	2968.6500	2	340	173	15.9	10.27	1224	380	2230	0.000	RoPr	-5.40	2.837	0.284	0.79	0.30	0.16	2.54	-25.18	3.550	20.12	-5.44	0.031	-5.425	0.167	0.79	0.018	0.790	0.096	3.59	-5.44	0.022	-5.425	0.120	0.79	0.012	0.790	0.069	3.60	117640	931	63.0	3	6.10	380	7.5	3	3.32	6127	0.047	15	0.049	4	C	C	A	19.9	938.5	50.5	463.3	Basel Landschaft, Amt für Geoinformation	-	Rust stains on the back of the 0-point excenter. / Strong dent in the rear of the housing (left side) @ ca. 423mm. / Cover for tensioning screw missing on top of the staff.	0387_20250312_0950_Leica GPCL3_80912 (A)_Basel Landschaft, Amt für Geoinformation.mat	Rust stains on the back of the 0-point excenter. Strong dent in the rear of the housing (left side) @ ca. 423mm. Cover for tensioning screw missing on top of the staff.
					-8.67	±3.436			1.11	±0.255

383	Leica GPCL3	80913 (B)	Basel Landschaft, Amt für Geoinformation	FALSE	3/6/2025	15:45:05	18:40:36	175.5	55.6875	2968.6500	2	340	173	15.9	10.26	1224	400	2250	0.000	RoPr	-10.91	3.186	0.341	5.43	0.45	0.20	3.06	-23.05	3.755	21.27	-10.97	0.039	-10.911	0.146	5.42	0.023	5.422	0.084	4.64	-10.96	0.028	-10.905	0.108	5.42	0.016	5.418	0.062	4.65	117640	625	62.9	3	6.36	285	7.4	3	3.26	6744	0.050	15	0.086	5	X	C	A	20.1	959.3	50.2	487.2	Basel Landschaft, Amt für Geoinformation	-	Slight rust stains on the back of the 0-point excenter	0383_20250306_1534_Leica GPCL3_80913 (B)_Basel Landschaft, Amt für Geoinformation.mat	Slight rust stains on the back of the 0-point excenter.
385	Leica GPCL3	80913 (B)	Basel Landschaft, Amt für Geoinformation	FALSE	3/7/2025	13:38:25	16:34:02	175.6	55.6875	2968.6500	2	340	173	15.9	10.27	1224	400	2250	0.000	RoPr	-16.98	2.567	0.340	5.62	0.38	0.20	3.06	-23.08	3.751	21.18	-17.02	0.035	-16.964	0.147	5.60	0.020	5.603	0.084	4.16	-17.02	0.025	-16.963	0.113	5.60	0.014	5.603	0.065	4.20	117640	685	62.9	3	6.34	361	7.4	3	3.26	6893	0.042	15	0.033	1	C	C	A	20.1	956.3	50.3	482.3	Basel Landschaft, Amt für Geoinformation	-	Rust stains on the back of the 0-point excenter.	0385_20250307_1326_Leica GPCL3_80913 (B)_Basel Landschaft, Amt für Geoinformation.mat	Rust stains on the back of the 0-point excenter.
386	Leica GPCL3	80913 (B)	Basel Landschaft, Amt für Geoinformation	FALSE	3/10/2025	11:30:24	14:26:41	176.3	55.6875	2968.6500	2	340	173	15.9	10.26	1224	400	2250	0.000	RoPr	-11.46	2.607	0.346	5.19	0.37	0.20	3.08	-23.13	3.759	21.21	-11.50	0.036	-11.465	0.146	5.17	0.020	5.201	0.084	4.18	-11.50	0.025	-11.464	0.105	5.17	0.014	5.205	0.060	4.18	117640	641	63.0	3	6.37	364	7.4	3	3.25	6989	0.044	15	0.039	8	C	C	A	20.0	944.5	50.4	446.1	Basel Landschaft, Amt für Geoinformation	-	Rust stains on the back of the 0-point excenter.	0386_20250310_1118_Leica GPCL3_80913 (B)_Basel Landschaft, Amt für Geoinformation.mat	Rust stains on the back of the 0-point excenter.
					-13.16	±2.737			5.40	±0.174